## Title
Varian Halcyon 1.0
6FFF Eclipse Data Tables
Revision 1, August 18, 2017

## Table of Contents
Table of contents
1. Data tables
Open field
Open Field Depth Doses
Open Field Profiles at 1.3 cm
Open Field Profiles at 5 cm
Open Field Profiles at 10 cm
Open Field Profiles at 20 cm
Open Field Profiles at 30 cm
Diagonal Profiles
Open Field Output Factors
2. Plots
Open Field

## Open Field Depth Dose
Percentage Depth Doses
Open field
SSD = 90 cm
Detector: IBA Dosimetry CC13 Ionization Chamber
Detector offset of -1.8mm (0.6*rcav) applied, interpolated to 1mm grid | [7 empty] | Lowess smoothed | [7 empty] | Percentage Diff - Original at 1mm grid vs Lowess Smoothed
Depth [cm] | 2x2cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 28x28cm2 | Depth [cm] | 2x2cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 28x28cm2 | Depth [cm] | 2x2cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 28x28cm2
0 | 58.55 | 58.9 | 59.34 | 60.72 | 62.24 | 67.8 | 70.4 | 0 | 58.14 | 58.9 | 59.21 | 60.75 | 62 | 67.49 | 70.31 | 0 | -0.7003% | 0% | -0.2191% | 0.04941% | -0.3856% | -0.4572% | -0.1278%
0.1 | 64.35 | 63.86 | 63.87 | 65.08 | 66.82 | 72.37 | 74.14 | 0.1 | 65.35 | 64.58 | 64.15 | 65.74 | 67.25 | 73.46 | 74.39 | 0.1 | 1.554% | 1.127% | 0.4384% | 1.014% | 0.6435% | 1.506% | 0.3372%
0.2 | 72.39 | 71.5 | 70.83 | 71.59 | 74.24 | 79.68 | 79.72 | 0.2 | 73.16 | 71.83 | 71.36 | 71.7 | 74.38 | 79.56 | 79.98 | 0.2 | 1.064% | 0.4615% | 0.7483% | 0.1537% | 0.1886% | -0.1506% | 0.3261%
0.3 | 81.6 | 80.25 | 80.19 | 80.67 | 82.28 | 86.23 | 86.4 | 0.3 | 81.24 | 80 | 79.53 | 79.46 | 81.94 | 85.92 | 86.19 | 0.3 | -0.4412% | -0.3115% | -0.823% | -1.5% | -0.4132% | -0.3595% | -0.2431%
0.4 | 87.59 | 86.46 | 86.12 | 86.47 | 88.05 | 90.76 | 91.13 | 0.4 | 87.37 | 86.31 | 86.08 | 86.74 | 87.74 | 90.66 | 90.85 | 0.4 | -0.2512% | -0.1735% | -0.04645% | 0.3122% | -0.3521% | -0.1102% | -0.3073%
0.5 | 92.05 | 90.8 | 90.69 | 91.07 | 91.87 | 93.95 | 94.29 | 0.5 | 91.8 | 90.85 | 90.44 | 90.86 | 91.74 | 94 | 94.21 | 0.5 | -0.2716% | 0.05507% | -0.2757% | -0.2306% | -0.1415% | 0.05322% | -0.08484%
0.6 | 94.91 | 93.95 | 93.53 | 94.12 | 94.7 | 96.19 | 96.52 | 0.6 | 94.68 | 93.92 | 93.52 | 94.12 | 94.65 | 96.1 | 96.36 | 0.6 | -0.2423% | -0.03193% | -0.01069% | 0% | -0.0528% | -0.09356% | -0.1658%
0.7 | 96.93 | 96.23 | 95.87 | 96.28 | 96.71 | 97.64 | 97.98 | 0.7 | 96.92 | 96.28 | 95.89 | 96.07 | 96.59 | 97.69 | 97.73 | 0.7 | -0.01032% | 0.05196% | 0.02086% | -0.2181% | -0.1241% | 0.05121% | -0.2552%
0.8 | 98.38 | 97.75 | 97.47 | 97.71 | 97.93 | 98.65 | 98.87 | 0.8 | 98.4 | 97.82 | 97.41 | 97.8 | 97.88 | 98.64 | 98.91 | 0.8 | 0.02033% | 0.07161% | -0.06156% | 0.09211% | -0.05106% | -0.01014% | 0.04046%
0.9 | 99.32 | 98.77 | 98.63 | 98.81 | 99 | 99.45 | 99.41 | 0.9 | 99.13 | 98.9 | 98.47 | 98.83 | 98.9 | 99.34 | 99.42 | 0.9 | -0.1913% | 0.1316% | -0.1622% | 0.02024% | -0.101% | -0.1106% | 0.01006%
1 | 99.75 | 99.5 | 99.25 | 99.44 | 99.51 | 99.71 | 99.85 | 1 | 99.75 | 99.57 | 99.26 | 99.48 | 99.49 | 99.77 | 99.81 | 1 | 0% | 0.07035% | 0.01008% | 0.04023% | -0.0201% | 0.06017% | -0.04006%
1.1 | 99.96 | 99.81 | 99.68 | 99.75 | 99.94 | 99.92 | 100 | 1.1 | 99.96 | 99.94 | 99.72 | 99.8 | 99.77 | 99.95 | 100 | 1.1 | 0% | 0.1302% | 0.04013% | 0.05013% | -0.1701% | 0.03002% | 0%
1.2 | 100 | 100 | 99.96 | 100 | 99.96 | 100 | 99.99 | 1.2 | 100 | 100 | 100 | 100 | 100 | 100 | 99.97 | 1.2 | 0% | 0% | 0.04002% | 0% | 0.04002% | 0% | -0.02%
1.3 | 99.83 | 99.78 | 100 | 99.9 | 100 | 99.78 | 99.77 | 1.3 | 99.8 | 99.94 | 99.98 | 100 | 99.92 | 99.89 | 99.76 | 1.3 | -0.03005% | 0.1604% | -0.02% | 0.1001% | -0.08% | 0.1102% | -0.01002%
1.4 | 99.43 | 99.65 | 99.81 | 99.81 | 99.83 | 99.59 | 99.54 | 1.4 | 99.45 | 99.76 | 99.83 | 99.88 | 99.8 | 99.67 | 99.5 | 1.4 | 0.02011% | 0.1104% | 0.02004% | 0.07013% | -0.03005% | 0.08033% | -0.04018%
1.5 | 99.12 | 99.42 | 99.54 | 99.57 | 99.67 | 99.28 | 99.19 | 1.5 | 99.06 | 99.55 | 99.59 | 99.64 | 99.6 | 99.4 | 99.35 | 1.5 | -0.06053% | 0.1308% | 0.05023% | 0.0703% | -0.07023% | 0.1209% | 0.1613%
1.6 | 98.63 | 99.05 | 99.19 | 99.14 | 99.26 | 98.95 | 99.06 | 1.6 | 98.58 | 99.21 | 99.27 | 99.28 | 99.28 | 99.07 | 99.08 | 1.6 | -0.05069% | 0.1615% | 0.08065% | 0.1412% | 0.02015% | 0.1213% | 0.02019%
1.7 | 97.92 | 98.68 | 98.85 | 98.97 | 98.96 | 98.78 | 98.68 | 1.7 | 97.99 | 98.77 | 98.83 | 98.95 | 98.94 | 98.75 | 98.72 | 1.7 | 0.07149% | 0.0912% | -0.02023% | -0.02021% | -0.02021% | -0.03037% | 0.04054%
1.8 | 97.47 | 98.06 | 98.36 | 98.42 | 98.53 | 98.23 | 98.34 | 1.8 | 97.44 | 98.29 | 98.38 | 98.54 | 98.57 | 98.37 | 98.35 | 1.8 | -0.03078% | 0.2346% | 0.02033% | 0.1219% | 0.0406% | 0.1425% | 0.01017%
1.9 | 96.88 | 97.71 | 97.91 | 97.96 | 98.12 | 97.87 | 97.95 | 1.9 | 96.89 | 97.85 | 97.94 | 98.09 | 98.15 | 97.95 | 97.92 | 1.9 | 0.01032% | 0.1433% | 0.03064% | 0.1327% | 0.03057% | 0.08174% | -0.03063%
2 | 96.23 | 97.19 | 97.36 | 97.69 | 97.59 | 97.46 | 97.65 | 2 | 96.3 | 97.35 | 97.48 | 97.7 | 97.63 | 97.56 | 97.51 | 2 | 0.07274% | 0.1646% | 0.1233% | 0.01024% | 0.04099% | 0.1026% | -0.1434%
2.1 | 95.72 | 96.63 | 96.99 | 97.1 | 97.31 | 97.06 | 97.18 | 2.1 | 95.68 | 96.76 | 97.05 | 97.23 | 97.17 | 97.16 | 97.15 | 2.1 | -0.04179% | 0.1345% | 0.06186% | 0.1339% | -0.1439% | 0.103% | -0.03087%
2.2 | 94.92 | 95.95 | 96.39 | 96.59 | 96.83 | 96.64 | 96.75 | 2.2 | 95.02 | 96.12 | 96.55 | 96.72 | 96.77 | 96.73 | 96.74 | 2.2 | 0.1054% | 0.1772% | 0.166% | 0.1346% | -0.06196% | 0.09313% | -0.01034%
2.3 | 94.43 | 95.35 | 95.97 | 96.1 | 96.32 | 96.2 | 96.3 | 2.3 | 94.36 | 95.51 | 96.01 | 96.24 | 96.31 | 96.28 | 96.32 | 2.3 | -0.07413% | 0.1678% | 0.04168% | 0.1457% | -0.01038% | 0.08316% | 0.02077%
2.4 | 93.58 | 94.77 | 95.31 | 95.64 | 95.74 | 95.8 | 95.95 | 2.4 | 93.66 | 94.94 | 95.43 | 95.7 | 95.85 | 95.91 | 95.98 | 2.4 | 0.08549% | 0.1794% | 0.1259% | 0.06274% | 0.1149% | 0.1148% | 0.03127%
2.5 | 93 | 94.18 | 94.88 | 95.24 | 95.38 | 95.4 | 95.6 | 2.5 | 93.02 | 94.39 | 94.92 | 95.22 | 95.42 | 95.5 | 95.61 | 2.5 | 0.02151% | 0.223% | 0.04216% | -0.021% | 0.04194% | 0.1048% | 0.01046%
2.6 | 92.42 | 93.67 | 94.32 | 94.73 | 94.9 | 94.9 | 95.2 | 2.6 | 92.4 | 93.84 | 94.43 | 94.81 | 94.94 | 95.02 | 95.21 | 2.6 | -0.02164% | 0.1815% | 0.1166% | 0.08445% | 0.04215% | 0.1264% | 0.0105%
2.7 | 91.65 | 93.13 | 93.9 | 94.16 | 94.4 | 94.5 | 94.8 | 2.7 | 91.72 | 93.29 | 93.89 | 94.26 | 94.44 | 94.61 | 94.8 | 2.7 | 0.07638% | 0.1718% | -0.01065% | 0.1062% | 0.04237% | 0.1164% | 0%
2.8 | 90.98 | 92.55 | 93.22 | 93.6 | 93.94 | 94 | 94.3 | 2.8 | 91.03 | 92.7 | 93.33 | 93.72 | 93.95 | 94.2 | 94.35 | 2.8 | 0.05496% | 0.1621% | 0.118% | 0.1282% | 0.01065% | 0.2128% | 0.05302%
2.9 | 90.38 | 91.95 | 92.79 | 93.16 | 93.41 | 93.7 | 94 | 2.9 | 90.4 | 92.13 | 92.81 | 93.22 | 93.45 | 93.79 | 93.94 | 2.9 | 0.02213% | 0.1958% | 0.02155% | 0.06441% | 0.04282% | 0.09605% | -0.06383%
3 | 89.8 | 91.41 | 92.24 | 92.6 | 92.97 | 93.3 | 93.4 | 3 | 89.73 | 91.53 | 92.27 | 92.7 | 92.98 | 93.39 | 93.52 | 3 | -0.07795% | 0.1313% | 0.03252% | 0.108% | 0.01076% | 0.09646% | 0.1285%
3.1 | 89.08 | 90.71 | 91.59 | 92.05 | 92.47 | 92.8 | 93.1 | 3.1 | 89.09 | 90.91 | 91.7 | 92.2 | 92.54 | 92.9 | 93.09 | 3.1 | 0.01123% | 0.2205% | 0.1201% | 0.163% | 0.0757% | 0.1078% | -0.01074%
3.2 | 88.44 | 90.11 | 91.09 | 91.61 | 92.06 | 92.3 | 92.6 | 3.2 | 88.48 | 90.32 | 91.11 | 91.7 | 92.07 | 92.41 | 92.65 | 3.2 | 0.04523% | 0.233% | 0.02196% | 0.09824% | 0.01086% | 0.1192% | 0.054%
3.3 | 87.81 | 89.61 | 90.61 | 91.05 | 91.56 | 91.9 | 92.2 | 3.3 | 87.82 | 89.75 | 90.55 | 91.21 | 91.56 | 92 | 92.23 | 3.3 | 0.01139% | 0.1562% | -0.06622% | 0.1757% | 0% | 0.1088% | 0.03254%
3.4 | 87.11 | 88.91 | 89.97 | 90.63 | 91.16 | 91.5 | 91.7 | 3.4 | 87.15 | 89.16 | 90 | 90.71 | 91.05 | 91.58 | 91.8 | 3.4 | 0.04592% | 0.2812% | 0.03334% | 0.08827% | -0.1207% | 0.08743% | 0.1091%
3.5 | 86.52 | 88.51 | 89.4 | 90.13 | 90.56 | 91.03 | 91.4 | 3.5 | 86.52 | 88.61 | 89.5 | 90.2 | 90.55 | 91.15 | 91.42 | 3.5 | 0% | 0.113% | 0.1119% | 0.07767% | -0.01104% | 0.1318% | 0.02188%
3.6 | 85.94 | 87.81 | 88.92 | 89.53 | 90.06 | 90.67 | 91 | 3.6 | 85.89 | 88 | 89.03 | 89.68 | 90.06 | 90.74 | 91 | 3.6 | -0.05818% | 0.2164% | 0.1237% | 0.1675% | 0% | 0.0772% | 0%
3.7 | 85.19 | 87.31 | 88.46 | 89.13 | 89.46 | 90.22 | 90.53 | 3.7 | 85.26 | 87.39 | 88.56 | 89.19 | 89.52 | 90.31 | 90.59 | 3.7 | 0.08217% | 0.09163% | 0.113% | 0.06732% | 0.06707% | 0.09976% | 0.06628%
3.8 | 84.69 | 86.71 | 87.96 | 88.43 | 89.06 | 89.68 | 90.16 | 3.8 | 84.69 | 86.82 | 88.06 | 88.63 | 89.02 | 89.82 | 90.21 | 3.8 | 0% | 0.1269% | 0.1137% | 0.2262% | -0.04491% | 0.1561% | 0.05546%
3.9 | 83.94 | 86.01 | 87.46 | 88.03 | 88.56 | 89.34 | 89.79 | 3.9 | 84.08 | 86.22 | 87.48 | 88.1 | 88.55 | 89.35 | 89.77 | 3.9 | 0.1668% | 0.2442% | 0.02287% | 0.07952% | -0.01129% | 0.01119% | -0.02227%
4 | 83.43 | 85.51 | 86.76 | 87.53 | 88.16 | 88.82 | 89.21 | 4 | 83.43 | 85.61 | 86.88 | 87.6 | 88.05 | 88.94 | 89.33 | 4 | 0% | 0.1169% | 0.1383% | 0.07997% | -0.1248% | 0.1351% | 0.1345%
4.1 | 82.8 | 84.91 | 86.26 | 86.93 | 87.56 | 88.29 | 88.95 | 4.1 | 82.77 | 85.03 | 86.34 | 87.1 | 87.56 | 88.49 | 88.91 | 4.1 | -0.03623% | 0.1413% | 0.09274% | 0.1956% | 0% | 0.2265% | -0.04497%
4.2 | 82.12 | 84.25 | 85.76 | 86.53 | 87 | 88.05 | 88.33 | 4.2 | 82.19 | 84.45 | 85.78 | 86.62 | 87.08 | 88.06 | 88.41 | 4.2 | 0.08524% | 0.2374% | 0.02332% | 0.104% | 0.09195% | 0.01136% | 0.09057%
4.3 | 81.69 | 83.77 | 85.16 | 86.03 | 86.53 | 87.51 | 87.97 | 4.3 | 81.65 | 83.89 | 85.24 | 86.11 | 86.62 | 87.67 | 87.98 | 4.3 | -0.04897% | 0.1432% | 0.09394% | 0.09299% | 0.104% | 0.1828% | 0.01137%
4.4 | 81.1 | 83.19 | 84.66 | 85.57 | 86.12 | 87.16 | 87.55 | 4.4 | 81.07 | 83.32 | 84.72 | 85.62 | 86.15 | 87.26 | 87.6 | 4.4 | -0.03699% | 0.1563% | 0.07087% | 0.05843% | 0.03484% | 0.1147% | 0.05711%
4.5 | 80.5 | 82.7 | 84.06 | 84.98 | 85.71 | 86.78 | 87.07 | 4.5 | 80.48 | 82.74 | 84.15 | 85.12 | 85.69 | 86.86 | 87.15 | 4.5 | -0.02484% | 0.04837% | 0.1071% | 0.1647% | -0.02333% | 0.09219% | 0.09188%
4.6 | 79.92 | 81.96 | 83.56 | 84.49 | 85.21 | 86.3 | 86.66 | 4.6 | 79.89 | 82.16 | 83.63 | 84.6 | 85.2 | 86.45 | 86.72 | 4.6 | -0.03754% | 0.244% | 0.08377% | 0.1302% | -0.01174% | 0.1738% | 0.06924%
4.7 | 79.2 | 81.55 | 83.11 | 83.98 | 84.7 | 85.86 | 86.26 | 4.7 | 79.25 | 81.68 | 83.12 | 84.09 | 84.7 | 86.01 | 86.33 | 4.7 | 0.06313% | 0.1594% | 0.01203% | 0.131% | 0% | 0.1747% | 0.08115%
4.8 | 78.62 | 81.08 | 82.4 | 83.51 | 84.19 | 85.51 | 85.92 | 4.8 | 78.67 | 81.18 | 82.6 | 83.59 | 84.22 | 85.56 | 85.95 | 4.8 | 0.0636% | 0.1233% | 0.2427% | 0.0958% | 0.03563% | 0.05847% | 0.03492%
4.9 | 78.2 | 80.54 | 82.02 | 82.99 | 83.73 | 84.99 | 85.61 | 4.9 | 78.13 | 80.6 | 82.1 | 83.11 | 83.74 | 85.07 | 85.56 | 4.9 | -0.08951% | 0.0745% | 0.09754% | 0.1446% | 0.01194% | 0.09413% | -0.0584%
5 | 77.51 | 79.85 | 81.62 | 82.57 | 83.29 | 84.55 | 85.06 | 5 | 77.53 | 80.04 | 81.62 | 82.65 | 83.25 | 84.62 | 85.13 | 5 | 0.0258% | 0.2379% | 0% | 0.09689% | -0.04802% | 0.08279% | 0.08229%
5.1 | 76.83 | 79.46 | 81.09 | 82 | 82.71 | 84.14 | 84.66 | 5.1 | 76.98 | 79.52 | 81.08 | 82.14 | 82.77 | 84.21 | 84.7 | 5.1 | 0.1952% | 0.07551% | -0.01233% | 0.1707% | 0.07254% | 0.08319% | 0.04725%
5.2 | 76.4 | 78.85 | 80.49 | 81.47 | 82.39 | 83.66 | 84.31 | 5.2 | 76.42 | 78.97 | 80.55 | 81.6 | 82.32 | 83.79 | 84.27 | 5.2 | 0.02618% | 0.1522% | 0.07454% | 0.1596% | -0.08496% | 0.1554% | -0.04744%
5.3 | 75.8 | 78.27 | 79.81 | 81.21 | 81.82 | 83.34 | 83.76 | 5.3 | 75.84 | 78.41 | 80.08 | 81.11 | 81.84 | 83.35 | 83.84 | 5.3 | 0.05277% | 0.1789% | 0.3383% | -0.1231% | 0.02444% | 0.012% | 0.09551%
5.4 | 75.36 | 77.73 | 79.62 | 80.5 | 81.29 | 82.79 | 83.42 | 5.4 | 75.33 | 77.85 | 79.6 | 80.62 | 81.34 | 82.87 | 83.45 | 5.4 | -0.03981% | 0.1544% | -0.02512% | 0.1491% | 0.06151% | 0.09663% | 0.03596%
5.5 | 74.76 | 77.2 | 78.9 | 80.06 | 80.87 | 82.33 | 83.09 | 5.5 | 74.79 | 77.29 | 79.02 | 80.13 | 80.93 | 82.45 | 83.08 | 5.5 | 0.04013% | 0.1166% | 0.1521% | 0.08743% | 0.07419% | 0.1458% | -0.01204%
5.6 | 74.26 | 76.63 | 78.41 | 79.57 | 80.57 | 81.91 | 82.62 | 5.6 | 74.24 | 76.76 | 78.5 | 79.64 | 80.49 | 82.05 | 82.65 | 5.6 | -0.02693% | 0.1696% | 0.1148% | 0.08797% | -0.09929% | 0.1709% | 0.03631%
5.7 | 73.56 | 76.2 | 77.82 | 79.02 | 79.92 | 81.61 | 82.19 | 5.7 | 73.68 | 76.29 | 78.01 | 79.15 | 79.99 | 81.64 | 82.2 | 5.7 | 0.1631% | 0.1181% | 0.2442% | 0.1645% | 0.08759% | 0.03676% | 0.01217%
5.8 | 73.16 | 75.61 | 77.47 | 78.6 | 79.55 | 81.16 | 81.78 | 5.8 | 73.14 | 75.77 | 77.51 | 78.64 | 79.53 | 81.24 | 81.81 | 5.8 | -0.02734% | 0.2116% | 0.05163% | 0.05089% | -0.02514% | 0.09857% | 0.03668%
5.9 | 72.46 | 75.07 | 76.91 | 78.13 | 79.09 | 80.75 | 81.48 | 5.9 | 72.63 | 75.21 | 76.99 | 78.17 | 79.09 | 80.84 | 81.45 | 5.9 | 0.2346% | 0.1865% | 0.104% | 0.0512% | 0% | 0.1115% | -0.03682%
6 | 72.06 | 74.62 | 76.44 | 77.61 | 78.5 | 80.33 | 81.02 | 6 | 72.08 | 74.68 | 76.48 | 77.69 | 78.58 | 80.4 | 81.04 | 6 | 0.02775% | 0.08041% | 0.05233% | 0.1031% | 0.1019% | 0.08714% | 0.02469%
6.1 | 71.36 | 74.05 | 75.86 | 77.12 | 78.05 | 79.86 | 80.5 | 6.1 | 71.51 | 74.15 | 75.96 | 77.19 | 78.04 | 79.96 | 80.58 | 6.1 | 0.2102% | 0.135% | 0.1318% | 0.09077% | -0.01281% | 0.1252% | 0.09938%
6.2 | 70.86 | 73.48 | 75.44 | 76.69 | 77.53 | 79.41 | 80.18 | 6.2 | 70.96 | 73.61 | 75.47 | 76.68 | 77.57 | 79.52 | 80.18 | 6.2 | 0.1411% | 0.1769% | 0.03977% | -0.01304% | 0.05159% | 0.1385% | 0%
6.3 | 70.46 | 72.97 | 74.86 | 76.17 | 77.13 | 79.02 | 79.8 | 6.3 | 70.43 | 73.13 | 74.96 | 76.17 | 77.15 | 79.1 | 79.77 | 6.3 | -0.04258% | 0.2193% | 0.1336% | 0% | 0.02593% | 0.1012% | -0.03759%
6.4 | 69.86 | 72.39 | 74.45 | 75.59 | 76.75 | 78.62 | 79.37 | 6.4 | 69.88 | 72.63 | 74.47 | 75.7 | 76.73 | 78.7 | 79.36 | 6.4 | 0.02863% | 0.3315% | 0.02686% | 0.1455% | -0.02606% | 0.1018% | -0.0126%
6.5 | 69.36 | 72.01 | 73.93 | 75.28 | 76.23 | 78.22 | 78.86 | 6.5 | 69.37 | 72.1 | 73.99 | 75.3 | 76.25 | 78.3 | 78.93 | 6.5 | 0.01442% | 0.125% | 0.08116% | 0.02657% | 0.02624% | 0.1023% | 0.08876%
6.6 | 68.86 | 71.42 | 73.43 | 74.75 | 75.71 | 77.78 | 78.48 | 6.6 | 68.87 | 71.57 | 73.48 | 74.83 | 75.78 | 77.88 | 78.5 | 6.6 | 0.01452% | 0.21% | 0.06809% | 0.107% | 0.09246% | 0.1286% | 0.02548%
6.7 | 68.36 | 70.97 | 72.85 | 74.2 | 75.39 | 77.45 | 78.22 | 6.7 | 68.38 | 71.07 | 73 | 74.35 | 75.35 | 77.47 | 78.12 | 6.7 | 0.02926% | 0.1409% | 0.2059% | 0.2022% | -0.05306% | 0.02582% | -0.1278%
6.8 | 67.86 | 70.49 | 72.55 | 73.93 | 74.89 | 76.96 | 77.71 | 6.8 | 67.93 | 70.6 | 72.56 | 73.92 | 74.92 | 77.05 | 77.77 | 6.8 | 0.1032% | 0.1561% | 0.01378% | -0.01353% | 0.04006% | 0.1169% | 0.07721%
6.9 | 67.32 | 70.03 | 72.03 | 73.34 | 74.46 | 76.51 | 77.37 | 6.9 | 67.45 | 70.14 | 72.08 | 73.46 | 74.5 | 76.61 | 77.41 | 6.9 | 0.1931% | 0.1571% | 0.06942% | 0.1636% | 0.05372% | 0.1307% | 0.0517%
7 | 66.92 | 69.47 | 71.53 | 72.96 | 74.04 | 76.07 | 76.89 | 7 | 66.95 | 69.64 | 71.59 | 73.01 | 74.04 | 76.19 | 77.03 | 7 | 0.04483% | 0.2447% | 0.08388% | 0.06853% | 0% | 0.1577% | 0.1821%
7.1 | 66.51 | 69.1 | 71.09 | 72.45 | 73.63 | 75.83 | 76.43 | 7.1 | 66.48 | 69.17 | 71.15 | 72.54 | 73.56 | 75.79 | 76.58 | 7.1 | -0.04511% | 0.1013% | 0.0844% | 0.1242% | -0.09507% | -0.05275% | 0.1963%
7.2 | 66 | 68.6 | 70.67 | 71.95 | 73.05 | 75.35 | 76.13 | 7.2 | 65.99 | 68.68 | 70.72 | 72.04 | 73.1 | 75.38 | 76.15 | 7.2 | -0.01515% | 0.1166% | 0.07075% | 0.1251% | 0.06845% | 0.03981% | 0.02627%
7.3 | 65.51 | 67.96 | 70.17 | 71.47 | 72.68 | 75.07 | 75.78 | 7.3 | 65.5 | 68.11 | 70.23 | 71.61 | 72.67 | 75.01 | 75.8 | 7.3 | -0.01526% | 0.2207% | 0.08551% | 0.1959% | -0.01376% | -0.07993% | 0.02639%
7.4 | 64.97 | 67.63 | 69.67 | 71.12 | 72.21 | 74.56 | 75.37 | 7.4 | 65.03 | 67.62 | 69.73 | 71.17 | 72.24 | 74.64 | 75.42 | 7.4 | 0.09235% | -0.01479% | 0.08612% | 0.0703% | 0.04155% | 0.1073% | 0.06634%
7.5 | 64.51 | 67.14 | 69.2 | 70.63 | 71.75 | 74.14 | 74.95 | 7.5 | 64.54 | 67.21 | 69.27 | 70.73 | 71.8 | 74.21 | 75 | 7.5 | 0.0465% | 0.1043% | 0.1012% | 0.1416% | 0.06969% | 0.09442% | 0.06671%
7.6 | 64 | 66.69 | 68.72 | 70.25 | 71.32 | 73.75 | 74.56 | 7.6 | 64.07 | 66.79 | 68.79 | 70.33 | 71.36 | 73.79 | 74.58 | 7.6 | 0.1094% | 0.1499% | 0.1019% | 0.1139% | 0.05609% | 0.05424% | 0.02682%
7.7 | 63.62 | 66.25 | 68.27 | 69.78 | 70.93 | 73.36 | 74.18 | 7.7 | 63.6 | 66.33 | 68.33 | 69.88 | 70.93 | 73.37 | 74.22 | 7.7 | -0.03144% | 0.1208% | 0.08789% | 0.1433% | 0% | 0.01363% | 0.05392%
7.8 | 63 | 65.64 | 67.88 | 69.27 | 70.55 | 72.91 | 73.88 | 7.8 | 63.11 | 65.87 | 67.9 | 69.41 | 70.49 | 72.94 | 73.89 | 7.8 | 0.1746% | 0.3504% | 0.02946% | 0.2021% | -0.08505% | 0.04115% | 0.01354%
7.9 | 62.72 | 65.29 | 67.37 | 69 | 69.95 | 72.46 | 73.45 | 7.9 | 62.69 | 65.37 | 67.43 | 69.02 | 70.03 | 72.56 | 73.49 | 7.9 | -0.04783% | 0.1225% | 0.08906% | 0.02899% | 0.1144% | 0.138% | 0.05446%
8 | 62.17 | 64.72 | 66.94 | 68.48 | 69.63 | 72.07 | 73.06 | 8 | 62.23 | 64.87 | 67.02 | 68.58 | 69.63 | 72.2 | 73.08 | 8 | 0.09651% | 0.2318% | 0.1195% | 0.146% | 0% | 0.1804% | 0.02737%
8.1 | 61.71 | 64.32 | 66.54 | 67.98 | 69.28 | 71.75 | 72.65 | 8.1 | 61.75 | 64.43 | 66.63 | 68.09 | 69.27 | 71.84 | 72.66 | 8.1 | 0.06482% | 0.171% | 0.1353% | 0.1618% | -0.01443% | 0.1254% | 0.01376%
8.2 | 61.29 | 63.92 | 66.08 | 67.6 | 68.85 | 71.47 | 72.31 | 8.2 | 61.31 | 63.99 | 66.18 | 67.66 | 68.86 | 71.51 | 72.28 | 8.2 | 0.03263% | 0.1095% | 0.1513% | 0.08876% | 0.01452% | 0.05597% | -0.04149%
8.3 | 60.95 | 63.36 | 65.64 | 67.03 | 68.4 | 71.1 | 71.93 | 8.3 | 60.9 | 63.51 | 65.68 | 67.22 | 68.43 | 71.14 | 71.94 | 8.3 | -0.08203% | 0.2367% | 0.06094% | 0.2835% | 0.04386% | 0.05626% | 0.0139%
8.4 | 60.47 | 62.94 | 65.12 | 66.65 | 68.08 | 70.6 | 71.57 | 8.4 | 60.48 | 63.06 | 65.2 | 66.75 | 68.02 | 70.74 | 71.57 | 8.4 | 0.01654% | 0.1907% | 0.1229% | 0.15% | -0.08813% | 0.1983% | 0%
8.5 | 59.96 | 62.67 | 64.8 | 66.37 | 67.62 | 70.3 | 71.23 | 8.5 | 60 | 62.67 | 64.78 | 66.38 | 67.61 | 70.33 | 71.25 | 8.5 | 0.06671% | 0% | -0.03086% | 0.01507% | -0.01479% | 0.04267% | 0.02808%
8.6 | 59.52 | 62.15 | 64.31 | 65.86 | 67.21 | 69.8 | 70.89 | 8.6 | 59.56 | 62.24 | 64.36 | 65.92 | 67.21 | 69.89 | 70.88 | 8.6 | 0.0672% | 0.1448% | 0.07775% | 0.0911% | 0% | 0.1289% | -0.01411%
8.7 | 59.17 | 61.67 | 63.9 | 65.57 | 66.7 | 69.5 | 70.47 | 8.7 | 59.16 | 61.78 | 63.96 | 65.47 | 66.75 | 69.55 | 70.49 | 8.7 | -0.0169% | 0.1784% | 0.0939% | -0.1525% | 0.07496% | 0.07194% | 0.02838%
8.8 | 58.67 | 61.27 | 63.5 | 65.01 | 66.29 | 69.1 | 70.05 | 8.8 | 58.69 | 61.36 | 63.55 | 65.08 | 66.3 | 69.18 | 70.08 | 8.8 | 0.03409% | 0.1469% | 0.07874% | 0.1077% | 0.01509% | 0.1158% | 0.04283%
8.9 | 58.25 | 60.78 | 63.09 | 64.62 | 65.92 | 68.7 | 69.7 | 8.9 | 58.23 | 60.93 | 63.11 | 64.69 | 65.91 | 68.77 | 69.72 | 8.9 | -0.03433% | 0.2468% | 0.0317% | 0.1083% | -0.01517% | 0.1019% | 0.02869%
9 | 57.78 | 60.35 | 62.61 | 64.17 | 65.51 | 68.3 | 69.37 | 9 | 57.82 | 60.49 | 62.65 | 64.25 | 65.52 | 68.38 | 69.39 | 9 | 0.06923% | 0.232% | 0.06389% | 0.1247% | 0.01526% | 0.1171% | 0.02883%
9.1 | 57.44 | 59.96 | 62.09 | 63.82 | 65.04 | 68 | 68.9 | 9.1 | 57.42 | 60.07 | 62.19 | 63.87 | 65.15 | 68.04 | 68.99 | 9.1 | -0.03482% | 0.1835% | 0.1611% | 0.07835% | 0.1691% | 0.05882% | 0.1306%
9.2 | 56.99 | 59.56 | 61.79 | 63.19 | 64.71 | 67.6 | 68.6 | 9.2 | 56.99 | 59.68 | 61.77 | 63.49 | 64.77 | 67.67 | 68.64 | 9.2 | 0% | 0.2015% | -0.03237% | 0.4748% | 0.09272% | 0.1036% | 0.05831%
9.3 | 56.49 | 59.21 | 61.31 | 62.92 | 64.29 | 67.2 | 68.3 | 9.3 | 56.56 | 59.3 | 61.36 | 63.05 | 64.34 | 67.28 | 68.29 | 9.3 | 0.1239% | 0.152% | 0.08155% | 0.2066% | 0.07777% | 0.119% | -0.01464%
9.4 | 56.28 | 58.77 | 60.9 | 62.46 | 63.88 | 66.9 | 67.9 | 9.4 | 56.21 | 58.86 | 60.95 | 62.61 | 63.9 | 66.93 | 67.89 | 9.4 | -0.1244% | 0.1531% | 0.0821% | 0.2402% | 0.03131% | 0.04484% | -0.01473%
9.5 | 55.84 | 58.38 | 60.5 | 62.22 | 63.51 | 66.5 | 67.5 | 9.5 | 55.82 | 58.44 | 60.56 | 62.22 | 63.5 | 66.55 | 67.52 | 9.5 | -0.03582% | 0.1028% | 0.09917% | 0% | -0.01575% | 0.07519% | 0.02963%
9.6 | 55.39 | 57.88 | 60.19 | 61.82 | 63.08 | 66.1 | 67.1 | 9.6 | 55.39 | 58.01 | 60.17 | 61.83 | 63.1 | 66.16 | 67.13 | 9.6 | 0% | 0.2246% | -0.03323% | 0.01618% | 0.03171% | 0.09077% | 0.04471%
9.7 | 54.9 | 57.45 | 59.7 | 61.32 | 62.72 | 65.8 | 66.7 | 9.7 | 54.96 | 57.58 | 59.75 | 61.41 | 62.7 | 65.84 | 66.75 | 9.7 | 0.1093% | 0.2263% | 0.08375% | 0.1468% | -0.03189% | 0.06079% | 0.07496%
9.8 | 54.57 | 57.08 | 59.38 | 61.02 | 62.3 | 65.2 | 66.4 | 9.8 | 54.56 | 57.2 | 59.29 | 61.02 | 62.32 | 65.46 | 66.4 | 9.8 | -0.01833% | 0.2102% | -0.1516% | 0% | 0.0321% | 0.3988% | 0%
9.9 | 54.13 | 56.7 | 58.84 | 60.62 | 61.9 | 64.9 | 66.1 | 9.9 | 54.14 | 56.81 | 58.87 | 60.62 | 61.92 | 65.05 | 66.09 | 9.9 | 0.01847% | 0.194% | 0.05099% | 0% | 0.03231% | 0.2311% | -0.01513%
10 | 53.8 | 56.3 | 58.49 | 60.12 | 61.5 | 64.6 | 65.6 | 10 | 53.77 | 56.39 | 58.53 | 60.2 | 61.52 | 64.67 | 65.69 | 10 | -0.05576% | 0.1599% | 0.06839% | 0.1331% | 0.03252% | 0.1084% | 0.1372%
10.1 | 53.37 | 55.8 | 58.12 | 59.72 | 61.12 | 64.3 | 65.2 | 10.1 | 53.41 | 55.98 | 58.16 | 59.81 | 61.16 | 64.28 | 65.35 | 10.1 | 0.07495% | 0.3226% | 0.06882% | 0.1507% | 0.06545% | -0.0311% | 0.2301%
10.2 | 53.01 | 55.5 | 57.69 | 59.32 | 60.82 | 63.8 | 65 | 10.2 | 53.05 | 55.59 | 57.76 | 59.43 | 60.82 | 63.91 | 65.02 | 10.2 | 0.07546% | 0.1622% | 0.1213% | 0.1854% | 0% | 0.1724% | 0.03077%
10.3 | 52.67 | 55 | 57.3 | 58.92 | 60.37 | 63.5 | 64.7 | 10.3 | 52.65 | 55.18 | 57.35 | 59.02 | 60.41 | 63.57 | 64.68 | 10.3 | -0.03797% | 0.3273% | 0.08726% | 0.1697% | 0.06626% | 0.1102% | -0.03091%
10.4 | 52.23 | 54.7 | 56.9 | 58.62 | 60.01 | 63.2 | 64.3 | 10.4 | 52.24 | 54.81 | 56.96 | 58.63 | 60.01 | 63.2 | 64.34 | 10.4 | 0.01915% | 0.2011% | 0.1054% | 0.01706% | 0% | 0% | 0.06221%
10.5 | 51.81 | 54.4 | 56.61 | 58.22 | 59.71 | 62.8 | 64 | 10.5 | 51.85 | 54.47 | 56.6 | 58.27 | 59.68 | 62.83 | 63.99 | 10.5 | 0.07721% | 0.1287% | -0.01766% | 0.08588% | -0.05024% | 0.04777% | -0.01562%
10.6 | 51.46 | 54 | 56.09 | 57.92 | 59.28 | 62.4 | 63.5 | 10.6 | 51.49 | 54.1 | 56.19 | 57.9 | 59.3 | 62.48 | 63.59 | 10.6 | 0.0583% | 0.1852% | 0.1783% | -0.03453% | 0.03374% | 0.1282% | 0.1417%
10.7 | 51.13 | 53.6 | 55.82 | 57.42 | 58.91 | 62.1 | 63.4 | 10.7 | 51.14 | 53.7 | 55.83 | 57.5 | 58.92 | 62.17 | 63.25 | 10.7 | 0.01956% | 0.1866% | 0.01791% | 0.1393% | 0.01698% | 0.1127% | -0.2366%
10.8 | 50.86 | 53.2 | 55.42 | 57.12 | 58.64 | 61.73 | 62.9 | 10.8 | 50.81 | 53.29 | 55.5 | 57.11 | 58.59 | 61.82 | 62.92 | 10.8 | -0.09831% | 0.1692% | 0.1444% | -0.01751% | -0.08527% | 0.1458% | 0.0318%
10.9 | 50.26 | 52.8 | 54.91 | 56.82 | 58.24 | 61.34 | 62.6 | 10.9 | 50.4 | 52.89 | 55.16 | 56.74 | 58.23 | 61.49 | 62.58 | 10.9 | 0.2786% | 0.1705% | 0.4553% | -0.1408% | -0.01717% | 0.2445% | -0.03195%
11 | 50.07 | 52.4 | 54.74 | 56.25 | 57.84 | 61.06 | 62.2 | 11 | 50.02 | 52.53 | 54.78 | 56.38 | 57.85 | 61.14 | 62.23 | 11 | -0.09986% | 0.2481% | 0.07307% | 0.2311% | 0.01729% | 0.131% | 0.04823%
11.1 | 49.58 | 52.1 | 54.34 | 55.97 | 57.44 | 60.76 | 61.9 | 11.1 | 49.67 | 52.18 | 54.38 | 56.02 | 57.49 | 60.78 | 61.9 | 11.1 | 0.1815% | 0.1536% | 0.07361% | 0.08933% | 0.08705% | 0.03292% | 0%
11.2 | 49.37 | 51.7 | 53.94 | 55.64 | 57.14 | 60.31 | 61.64 | 11.2 | 49.35 | 51.8 | 54.02 | 55.68 | 57.15 | 60.42 | 61.61 | 11.2 | -0.04051% | 0.1934% | 0.1483% | 0.07189% | 0.0175% | 0.1824% | -0.04867%
11.3 | 49.01 | 51.4 | 53.64 | 55.26 | 56.74 | 60.04 | 61.33 | 11.3 | 49 | 51.43 | 53.65 | 55.32 | 56.73 | 60.12 | 61.28 | 11.3 | -0.0204% | 0.05837% | 0.01864% | 0.1086% | -0.01762% | 0.1332% | -0.08153%
11.4 | 48.58 | 51 | 53.24 | 54.85 | 56.34 | 59.75 | 60.92 | 11.4 | 48.64 | 51.09 | 53.27 | 54.95 | 56.34 | 59.8 | 60.93 | 11.4 | 0.1235% | 0.1765% | 0.05635% | 0.1823% | 0% | 0.08368% | 0.01641%
11.5 | 48.25 | 50.6 | 52.94 | 54.56 | 56.04 | 59.43 | 60.53 | 11.5 | 48.29 | 50.75 | 52.91 | 54.61 | 56.02 | 59.47 | 60.58 | 11.5 | 0.0829% | 0.2964% | -0.05667% | 0.09164% | -0.03569% | 0.06731% | 0.0826%
11.6 | 47.97 | 50.26 | 52.54 | 54.16 | 55.64 | 59.12 | 60.15 | 11.6 | 47.95 | 50.38 | 52.53 | 54.22 | 55.66 | 59.16 | 60.27 | 11.6 | -0.04169% | 0.2388% | -0.01903% | 0.1108% | 0.03595% | 0.06766% | 0.1995%
11.7 | 47.6 | 50 | 52.14 | 53.75 | 55.34 | 58.73 | 59.94 | 11.7 | 47.62 | 50.06 | 52.16 | 53.85 | 55.32 | 58.83 | 59.97 | 11.7 | 0.04202% | 0.12% | 0.03836% | 0.186% | -0.03614% | 0.1703% | 0.05005%
11.8 | 47.25 | 49.55 | 51.74 | 53.46 | 54.94 | 58.26 | 59.65 | 11.8 | 47.28 | 49.7 | 51.82 | 53.54 | 54.96 | 58.48 | 59.64 | 11.8 | 0.06349% | 0.3027% | 0.1546% | 0.1496% | 0.0364% | 0.3776% | -0.01676%
11.9 | 46.97 | 49.23 | 51.54 | 53.16 | 54.64 | 58.07 | 59.23 | 11.9 | 46.95 | 49.34 | 51.53 | 53.21 | 54.61 | 58.11 | 59.28 | 11.9 | -0.04258% | 0.2234% | -0.0194% | 0.09406% | -0.0549% | 0.06888% | 0.08442%
12 | 46.61 | 48.96 | 51.14 | 52.75 | 54.24 | 57.71 | 58.85 | 12 | 46.6 | 49.03 | 51.18 | 52.84 | 54.26 | 57.77 | 58.97 | 12 | -0.02145% | 0.143% | 0.07822% | 0.1706% | 0.03687% | 0.104% | 0.2039%
12.1 | 46.19 | 48.67 | 50.74 | 52.37 | 53.94 | 57.34 | 58.66 | 12.1 | 46.24 | 48.72 | 50.78 | 52.48 | 53.91 | 57.4 | 58.65 | 12.1 | 0.1082% | 0.1027% | 0.07883% | 0.21% | -0.05562% | 0.1046% | -0.01705%
12.2 | 45.89 | 48.21 | 50.44 | 52.06 | 53.54 | 56.93 | 58.33 | 12.2 | 45.92 | 48.33 | 50.41 | 52.14 | 53.55 | 57.03 | 58.32 | 12.2 | 0.06537% | 0.2489% | -0.05948% | 0.1537% | 0.01868% | 0.1757% | -0.01714%
12.3 | 45.63 | 47.94 | 50.03 | 51.76 | 53.14 | 56.69 | 58.04 | 12.3 | 45.64 | 48 | 50.06 | 51.8 | 53.2 | 56.71 | 58 | 12.3 | 0.02192% | 0.1252% | 0.05996% | 0.07728% | 0.1129% | 0.03528% | -0.06892%
12.4 | 45.21 | 47.6 | 49.73 | 51.35 | 52.94 | 56.41 | 57.72 | 12.4 | 45.33 | 47.69 | 49.75 | 51.45 | 52.92 | 56.4 | 57.66 | 12.4 | 0.2654% | 0.1891% | 0.04022% | 0.1947% | -0.03778% | -0.01773% | -0.104%
12.5 | 44.96 | 47.15 | 49.43 | 51.06 | 52.54 | 56.02 | 57.32 | 12.5 | 44.98 | 47.31 | 49.44 | 51.13 | 52.56 | 56.12 | 57.33 | 12.5 | 0.04448% | 0.3393% | 0.02023% | 0.1371% | 0.03807% | 0.1785% | 0.01745%
12.6 | 44.59 | 46.83 | 49.03 | 50.76 | 52.34 | 55.7 | 57.01 | 12.6 | 44.62 | 46.93 | 49.11 | 50.8 | 52.2 | 55.85 | 57.03 | 12.6 | 0.06728% | 0.2135% | 0.1632% | 0.0788% | -0.2675% | 0.2693% | 0.03508%
12.7 | 44.32 | 46.56 | 48.73 | 50.36 | 51.84 | 55.5 | 56.66 | 12.7 | 44.32 | 46.64 | 48.8 | 50.42 | 51.87 | 55.53 | 56.72 | 12.7 | 0% | 0.1718% | 0.1436% | 0.1191% | 0.05787% | 0.05405% | 0.1059%
12.8 | 44.01 | 46.33 | 48.48 | 50.04 | 51.64 | 55.08 | 56.43 | 12.8 | 44.04 | 46.37 | 48.47 | 50.11 | 51.59 | 55.17 | 56.4 | 12.8 | 0.06817% | 0.08634% | -0.02063% | 0.1399% | -0.09682% | 0.1634% | -0.05316%
12.9 | 43.72 | 45.84 | 48.03 | 49.85 | 51.27 | 54.85 | 56.04 | 12.9 | 43.73 | 46.07 | 48.1 | 49.81 | 51.28 | 54.87 | 56.09 | 12.9 | 0.02287% | 0.5017% | 0.1457% | -0.08024% | 0.0195% | 0.03646% | 0.08922%
13 | 43.3 | 45.66 | 47.75 | 49.33 | 50.87 | 54.47 | 55.83 | 13 | 43.39 | 45.72 | 47.79 | 49.44 | 50.92 | 54.56 | 55.8 | 13 | 0.2079% | 0.1314% | 0.08377% | 0.223% | 0.09829% | 0.1652% | -0.05373%
13.1 | 43.1 | 45.33 | 47.39 | 49.14 | 50.66 | 54.18 | 55.43 | 13.1 | 43.08 | 45.38 | 47.48 | 49.14 | 50.63 | 54.24 | 55.48 | 13.1 | -0.0464% | 0.1103% | 0.1899% | 0% | -0.05922% | 0.1107% | 0.0902%
13.2 | 42.8 | 44.94 | 47.15 | 48.87 | 50.28 | 53.89 | 55.14 | 13.2 | 42.79 | 45.06 | 47.19 | 48.88 | 50.29 | 53.92 | 55.18 | 13.2 | -0.02336% | 0.267% | 0.08484% | 0.02046% | 0.01989% | 0.05567% | 0.07254%
13.3 | 42.52 | 44.74 | 46.88 | 48.5 | 49.95 | 53.55 | 54.94 | 13.3 | 42.52 | 44.74 | 46.89 | 48.55 | 49.98 | 53.59 | 54.88 | 13.3 | 0% | 0% | 0.02133% | 0.1031% | 0.06006% | 0.0747% | -0.1092%
13.4 | 42.21 | 44.3 | 46.47 | 48.24 | 49.67 | 53.25 | 54.53 | 13.4 | 42.24 | 44.41 | 46.52 | 48.2 | 49.71 | 53.29 | 54.58 | 13.4 | 0.07107% | 0.2483% | 0.1076% | -0.08292% | 0.08053% | 0.07512% | 0.09169%
13.5 | 41.89 | 44.12 | 46.15 | 47.85 | 49.46 | 53.01 | 54.28 | 13.5 | 41.91 | 44.12 | 46.2 | 47.9 | 49.43 | 52.97 | 54.3 | 13.5 | 0.04774% | 0% | 0.1083% | 0.1045% | -0.06066% | -0.07546% | 0.03685%
13.6 | 41.62 | 43.72 | 45.88 | 47.56 | 49.08 | 52.55 | 53.98 | 13.6 | 41.6 | 43.82 | 45.9 | 47.6 | 49.09 | 52.68 | 54 | 13.6 | -0.04805% | 0.2287% | 0.04359% | 0.0841% | 0.02037% | 0.2474% | 0.03705%
13.7 | 41.2 | 43.48 | 45.56 | 47.18 | 48.76 | 52.47 | 53.68 | 13.7 | 41.28 | 43.56 | 45.57 | 47.27 | 48.74 | 52.4 | 53.7 | 13.7 | 0.1942% | 0.184% | 0.02195% | 0.1908% | -0.04102% | -0.1334% | 0.03726%
13.8 | 41 | 43.19 | 45.25 | 46.9 | 48.36 | 51.96 | 53.38 | 13.8 | 40.98 | 43.24 | 45.3 | 46.96 | 48.44 | 52.06 | 53.4 | 13.8 | -0.04878% | 0.1158% | 0.1105% | 0.1279% | 0.1654% | 0.1925% | 0.03747%
13.9 | 40.72 | 42.78 | 44.98 | 46.67 | 48.19 | 51.68 | 53.09 | 13.9 | 40.72 | 42.88 | 45 | 46.68 | 48.14 | 51.74 | 53.08 | 13.9 | 0% | 0.2338% | 0.04446% | 0.02143% | -0.1038% | 0.1161% | -0.01884%
14 | 40.41 | 42.63 | 44.66 | 46.3 | 47.75 | 51.46 | 52.73 | 14 | 40.44 | 42.58 | 44.67 | 46.35 | 47.81 | 51.47 | 52.77 | 14 | 0.07424% | -0.1173% | 0.02239% | 0.108% | 0.1257% | 0.01943% | 0.07586%
14.1 | 40.1 | 42.24 | 44.34 | 45.96 | 47.56 | 51.14 | 52.49 | 14.1 | 40.15 | 42.3 | 44.39 | 46.01 | 47.53 | 51.19 | 52.48 | 14.1 | 0.1247% | 0.142% | 0.1128% | 0.1088% | -0.06308% | 0.09777% | -0.01905%
14.2 | 39.9 | 42.02 | 44.08 | 45.64 | 47.17 | 50.83 | 52.21 | 14.2 | 39.87 | 42 | 44.1 | 45.72 | 47.22 | 50.91 | 52.21 | 14.2 | -0.07519% | -0.0476% | 0.04537% | 0.1753% | 0.106% | 0.1574% | 0%
14.3 | 39.53 | 41.62 | 43.75 | 45.33 | 46.96 | 50.59 | 51.94 | 14.3 | 39.59 | 41.72 | 43.76 | 45.4 | 46.93 | 50.62 | 51.95 | 14.3 | 0.1518% | 0.2403% | 0.02286% | 0.1544% | -0.06388% | 0.0593% | 0.01925%
14.4 | 39.29 | 41.38 | 43.36 | 45.06 | 46.55 | 50.31 | 51.64 | 14.4 | 39.3 | 41.46 | 43.44 | 45.11 | 46.6 | 50.34 | 51.65 | 14.4 | 0.02545% | 0.1933% | 0.1845% | 0.111% | 0.1074% | 0.05963% | 0.01936%
14.5 | 39.02 | 41.08 | 43.15 | 44.79 | 46.3 | 49.91 | 51.34 | 14.5 | 39.04 | 41.16 | 43.18 | 44.85 | 46.26 | 50.08 | 51.35 | 14.5 | 0.05126% | 0.1947% | 0.06952% | 0.134% | -0.08639% | 0.3406% | 0.01948%
14.6 | 38.83 | 40.84 | 42.86 | 44.51 | 45.95 | 49.69 | 51.03 | 14.6 | 38.8 | 40.91 | 42.92 | 44.54 | 45.98 | 49.79 | 51.07 | 14.6 | -0.07726% | 0.1714% | 0.14% | 0.0674% | 0.06529% | 0.2012% | 0.07839%
14.7 | 38.53 | 40.58 | 42.56 | 44.22 | 45.76 | 49.44 | 50.78 | 14.7 | 38.55 | 40.65 | 42.62 | 44.26 | 45.72 | 49.48 | 50.79 | 14.7 | 0.05191% | 0.1725% | 0.141% | 0.09046% | -0.08741% | 0.08091% | 0.01969%
14.8 | 37.93 | 40.29 | 42.26 | 43.95 | 45.36 | 49.17 | 50.48 | 14.8 | 38.24 | 40.34 | 42.31 | 44 | 45.43 | 49.19 | 50.5 | 14.8 | 0.8173% | 0.1241% | 0.1183% | 0.1138% | 0.1543% | 0.04068% | 0.03962%
14.9 | 37.93 | 39.94 | 42.06 | 43.67 | 45.19 | 48.81 | 50.26 | 14.9 | 37.96 | 40.03 | 42.03 | 43.72 | 45.16 | 48.89 | 50.23 | 14.9 | 0.07909% | 0.2253% | -0.07133% | 0.1145% | -0.06639% | 0.1639% | -0.05969%
15 | 37.73 | 39.77 | 41.78 | 43.36 | 44.87 | 48.56 | 49.95 | 15 | 37.76 | 39.8 | 41.78 | 43.43 | 44.88 | 48.61 | 49.95 | 15 | 0.07951% | 0.07543% | 0% | 0.1614% | 0.02229% | 0.103% | 0%
15.1 | 37.53 | 39.39 | 41.39 | 43.11 | 44.56 | 48.33 | 49.64 | 15.1 | 37.5 | 39.54 | 41.51 | 43.16 | 44.6 | 48.34 | 49.65 | 15.1 | -0.07994% | 0.3808% | 0.2899% | 0.116% | 0.08977% | 0.02069% | 0.02015%
15.2 | 37.23 | 39.25 | 41.29 | 42.84 | 44.31 | 47.95 | 49.35 | 15.2 | 37.21 | 39.29 | 41.27 | 42.88 | 44.32 | 48.04 | 49.38 | 15.2 | -0.05372% | 0.1019% | -0.04844% | 0.09337% | 0.02257% | 0.1877% | 0.06079%
15.3 | 36.93 | 38.96 | 40.9 | 42.54 | 44.01 | 47.73 | 49.18 | 15.3 | 36.94 | 39 | 40.97 | 42.58 | 44.03 | 47.74 | 49.15 | 15.3 | 0.02708% | 0.1027% | 0.1711% | 0.09403% | 0.04544% | 0.02095% | -0.061%
15.4 | 36.63 | 38.57 | 40.67 | 42.24 | 43.71 | 47.35 | 48.78 | 15.4 | 36.67 | 38.68 | 40.7 | 42.28 | 43.73 | 47.44 | 48.91 | 15.4 | 0.1092% | 0.2852% | 0.07376% | 0.0947% | 0.04576% | 0.1901% | 0.2665%
15.5 | 36.43 | 38.34 | 40.41 | 41.98 | 43.4 | 47.16 | 48.64 | 15.5 | 36.42 | 38.41 | 40.45 | 42.03 | 43.46 | 47.21 | 48.62 | 15.5 | -0.02745% | 0.1826% | 0.09899% | 0.1191% | 0.1382% | 0.106% | -0.04112%
15.6 | 36.23 | 38.12 | 40.17 | 41.74 | 43.29 | 46.92 | 48.18 | 15.6 | 36.18 | 38.18 | 40.2 | 41.77 | 43.24 | 46.96 | 48.32 | 15.6 | -0.138% | 0.1574% | 0.07468% | 0.07187% | -0.1155% | 0.08525% | 0.2906%
15.7 | 35.93 | 37.84 | 39.92 | 41.44 | 42.9 | 46.67 | 48.02 | 15.7 | 35.92 | 37.91 | 39.92 | 41.48 | 42.94 | 46.69 | 48.02 | 15.7 | -0.02783% | 0.185% | 0% | 0.09653% | 0.09324% | 0.04285% | 0%
15.8 | 35.73 | 37.7 | 39.56 | 41.18 | 42.61 | 46.37 | 47.71 | 15.8 | 35.68 | 37.66 | 39.62 | 41.22 | 42.63 | 46.42 | 47.72 | 15.8 | -0.1399% | -0.1061% | 0.1517% | 0.09713% | 0.04694% | 0.1078% | 0.02096%
15.9 | 35.43 | 37.37 | 39.38 | 40.93 | 42.37 | 46.04 | 47.5 | 15.9 | 35.45 | 37.43 | 39.37 | 40.99 | 42.38 | 46.15 | 47.43 | 15.9 | 0.05645% | 0.1606% | -0.02539% | 0.1466% | 0.0236% | 0.2389% | -0.1474%
16 | 35.23 | 37.17 | 39.07 | 40.76 | 42.18 | 45.85 | 47.18 | 16 | 35.23 | 37.21 | 39.09 | 40.73 | 42.15 | 45.88 | 47.16 | 16 | 0% | 0.1076% | 0.05119% | -0.0736% | -0.07112% | 0.06543% | -0.04239%
16.1 | 34.93 | 36.86 | 38.84 | 40.37 | 41.8 | 45.57 | 46.85 | 16.1 | 34.96 | 36.93 | 38.79 | 40.45 | 41.84 | 45.59 | 46.9 | 16.1 | 0.08589% | 0.1899% | -0.1287% | 0.1982% | 0.09569% | 0.04389% | 0.1067%
16.2 | 34.63 | 36.55 | 38.45 | 40.18 | 41.5 | 45.29 | 46.66 | 16.2 | 34.67 | 36.65 | 38.51 | 40.2 | 41.55 | 45.32 | 46.67 | 16.2 | 0.1155% | 0.2736% | 0.156% | 0.04978% | 0.1205% | 0.06624% | 0.02143%
16.3 | 34.43 | 36.37 | 38.27 | 39.86 | 41.33 | 44.98 | 46.47 | 16.3 | 34.44 | 36.41 | 38.3 | 39.94 | 41.28 | 45.04 | 46.45 | 16.3 | 0.02904% | 0.11% | 0.07839% | 0.2007% | -0.121% | 0.1334% | -0.04304%
16.4 | 34.26 | 36.1 | 38.08 | 39.64 | 40.96 | 44.69 | 46.16 | 16.4 | 34.22 | 36.18 | 38.08 | 39.65 | 41 | 44.78 | 46.18 | 16.4 | -0.1168% | 0.2216% | 0% | 0.02523% | 0.09766% | 0.2014% | 0.04333%
16.5 | 33.95 | 35.86 | 37.83 | 39.27 | 40.77 | 44.47 | 45.85 | 16.5 | 33.96 | 35.93 | 37.84 | 39.35 | 40.77 | 44.55 | 45.89 | 16.5 | 0.02946% | 0.1952% | 0.02643% | 0.2037% | 0% | 0.1799% | 0.08724%
16.6 | 33.74 | 35.6 | 37.48 | 39.07 | 40.51 | 44.26 | 45.61 | 16.6 | 33.75 | 35.67 | 37.58 | 39.12 | 40.53 | 44.28 | 45.64 | 16.6 | 0.02964% | 0.1966% | 0.2668% | 0.128% | 0.04937% | 0.04519% | 0.06578%
16.7 | 33.46 | 35.3 | 37.33 | 38.87 | 40.29 | 43.99 | 45.43 | 16.7 | 33.53 | 35.43 | 37.31 | 38.92 | 40.27 | 44.02 | 45.4 | 16.7 | 0.2092% | 0.3683% | -0.05358% | 0.1286% | -0.04964% | 0.0682% | -0.06604%
16.8 | 33.36 | 35.18 | 37.03 | 38.64 | 40.03 | 43.67 | 45.12 | 16.8 | 33.31 | 35.21 | 37.06 | 38.67 | 40.04 | 43.76 | 45.16 | 16.8 | -0.1499% | 0.08528% | 0.08102% | 0.07764% | 0.02498% | 0.2061% | 0.08865%
16.9 | 33.05 | 34.97 | 36.77 | 38.39 | 39.76 | 43.49 | 44.98 | 16.9 | 33.06 | 35 | 36.83 | 38.4 | 39.8 | 43.51 | 44.94 | 16.9 | 0.03026% | 0.08579% | 0.1632% | 0.02605% | 0.1006% | 0.04599% | -0.08893%
17 | 32.84 | 34.66 | 36.58 | 38.06 | 39.54 | 43.19 | 44.61 | 17 | 32.85 | 34.73 | 36.61 | 38.14 | 39.56 | 43.28 | 44.65 | 17 | 0.03045% | 0.202% | 0.08201% | 0.2102% | 0.05058% | 0.2084% | 0.08967%
17.1 | 32.66 | 34.4 | 36.35 | 37.84 | 39.36 | 42.95 | 44.39 | 17.1 | 32.62 | 34.47 | 36.37 | 37.87 | 39.32 | 43 | 44.38 | 17.1 | -0.1225% | 0.2035% | 0.05502% | 0.07928% | -0.1016% | 0.1164% | -0.02253%
17.2 | 32.34 | 34.15 | 36.09 | 37.58 | 39.05 | 42.68 | 44.08 | 17.2 | 32.37 | 34.24 | 36.12 | 37.63 | 39.05 | 42.73 | 44.16 | 17.2 | 0.09276% | 0.2635% | 0.08313% | 0.133% | 0% | 0.1172% | 0.1815%
17.3 | 32.15 | 33.97 | 35.89 | 37.4 | 38.8 | 42.49 | 43.97 | 17.3 | 32.15 | 34.01 | 35.89 | 37.42 | 38.79 | 42.52 | 43.93 | 17.3 | 0% | 0.1178% | 0% | 0.05348% | -0.02577% | 0.0706% | -0.09097%
17.4 | 31.95 | 33.69 | 35.57 | 37.09 | 38.53 | 42.19 | 43.6 | 17.4 | 31.95 | 33.8 | 35.64 | 37.17 | 38.55 | 42.31 | 43.66 | 17.4 | 0% | 0.3265% | 0.1968% | 0.2157% | 0.05191% | 0.2844% | 0.1376%
17.5 | 31.75 | 33.52 | 35.39 | 36.89 | 38.32 | 42.04 | 43.39 | 17.5 | 31.75 | 33.57 | 35.39 | 36.93 | 38.33 | 42.09 | 43.46 | 17.5 | 0% | 0.1492% | 0% | 0.1084% | 0.0261% | 0.1189% | 0.1613%
17.6 | 31.54 | 33.27 | 35.07 | 36.68 | 38.02 | 41.78 | 43.26 | 17.6 | 31.55 | 33.33 | 35.14 | 36.69 | 38.08 | 41.8 | 43.24 | 17.6 | 0.03171% | 0.1803% | 0.1996% | 0.02726% | 0.1578% | 0.04787% | -0.04623%
17.7 | 31.25 | 33.08 | 34.89 | 36.4 | 37.86 | 41.48 | 42.96 | 17.7 | 31.34 | 33.13 | 34.89 | 36.44 | 37.83 | 41.53 | 42.97 | 17.7 | 0.288% | 0.1511% | 0% | 0.1099% | -0.07924% | 0.1205% | 0.02328%
17.8 | 31.16 | 32.87 | 34.58 | 36.19 | 37.59 | 41.27 | 42.78 | 17.8 | 31.12 | 32.93 | 34.64 | 36.23 | 37.6 | 41.28 | 42.72 | 17.8 | -0.1284% | 0.1825% | 0.1735% | 0.1105% | 0.0266% | 0.02423% | -0.1403%
17.9 | 30.83 | 32.67 | 34.43 | 35.99 | 37.4 | 40.99 | 42.49 | 17.9 | 30.88 | 32.7 | 34.43 | 36.03 | 37.38 | 41.04 | 42.5 | 17.9 | 0.1622% | 0.09183% | 0% | 0.1111% | -0.05348% | 0.122% | 0.02353%
18 | 30.67 | 32.38 | 34.12 | 35.8 | 37.12 | 40.78 | 42.29 | 18 | 30.67 | 32.44 | 34.21 | 35.82 | 37.15 | 40.83 | 42.28 | 18 | 0% | 0.1853% | 0.2638% | 0.05587% | 0.08082% | 0.1226% | -0.02365%
18.1 | 30.45 | 32.18 | 34 | 35.48 | 36.87 | 40.57 | 41.97 | 18.1 | 30.46 | 32.24 | 33.98 | 35.56 | 36.92 | 40.59 | 42.02 | 18.1 | 0.03284% | 0.1865% | -0.05882% | 0.2255% | 0.1356% | 0.0493% | 0.1191%
18.2 | 30.34 | 31.98 | 33.67 | 35.28 | 36.69 | 40.3 | 41.78 | 18.2 | 30.28 | 32.04 | 33.76 | 35.29 | 36.67 | 40.35 | 41.79 | 18.2 | -0.1978% | 0.1876% | 0.2673% | 0.02834% | -0.05451% | 0.1241% | 0.02393%
18.3 | 30.13 | 31.78 | 33.6 | 35 | 36.43 | 40.1 | 41.57 | 18.3 | 30.09 | 31.84 | 33.57 | 35.04 | 36.44 | 40.12 | 41.55 | 18.3 | -0.1328% | 0.1888% | -0.08929% | 0.1143% | 0.02745% | 0.04988% | -0.04811%
18.4 | 29.84 | 31.58 | 33.26 | 34.79 | 36.19 | 39.9 | 41.28 | 18.4 | 29.86 | 31.64 | 33.34 | 34.83 | 36.23 | 39.89 | 41.3 | 18.4 | 0.06702% | 0.19% | 0.2405% | 0.115% | 0.1105% | -0.02506% | 0.04845%
18.5 | 29.64 | 31.38 | 33.11 | 34.59 | 35.99 | 39.6 | 41.09 | 18.5 | 29.65 | 31.44 | 33.11 | 34.63 | 35.97 | 39.66 | 41.08 | 18.5 | 0.03374% | 0.1912% | 0% | 0.1156% | -0.05557% | 0.1515% | -0.02434%
18.6 | 29.44 | 31.17 | 32.91 | 34.38 | 35.7 | 39.4 | 40.77 | 18.6 | 29.45 | 31.19 | 32.91 | 34.39 | 35.73 | 39.41 | 40.83 | 18.6 | 0.03397% | 0.06416% | 0% | 0.02909% | 0.08403% | 0.02538% | 0.1472%
18.7 | 29.24 | 30.91 | 32.68 | 34.1 | 35.6 | 39.1 | 40.68 | 18.7 | 29.25 | 30.96 | 32.72 | 34.14 | 35.53 | 39.15 | 40.59 | 18.7 | 0.0342% | 0.1618% | 0.1224% | 0.1173% | -0.1966% | 0.1279% | -0.2212%
18.8 | 29.04 | 30.67 | 32.48 | 33.91 | 35.31 | 38.9 | 40.28 | 18.8 | 29.05 | 30.76 | 32.5 | 33.95 | 35.32 | 38.94 | 40.33 | 18.8 | 0.03444% | 0.2934% | 0.06158% | 0.118% | 0.02832% | 0.1028% | 0.1241%
18.9 | 28.84 | 30.57 | 32.3 | 33.68 | 35.11 | 38.7 | 40.08 | 18.9 | 28.85 | 30.55 | 32.26 | 33.75 | 35.1 | 38.7 | 40.09 | 18.9 | 0.03467% | -0.06542% | -0.1238% | 0.2078% | -0.02848% | 0% | 0.02495%
19 | 28.64 | 30.29 | 32.01 | 33.49 | 34.8 | 38.4 | 39.88 | 19 | 28.65 | 30.34 | 32.03 | 33.53 | 34.84 | 38.45 | 39.89 | 19 | 0.03492% | 0.1651% | 0.06248% | 0.1194% | 0.1149% | 0.1302% | 0.02508%
19.1 | 28.45 | 30.1 | 31.8 | 33.28 | 34.6 | 38.2 | 39.7 | 19.1 | 28.46 | 30.15 | 31.83 | 33.29 | 34.61 | 38.24 | 39.7 | 19.1 | 0.03515% | 0.1661% | 0.09434% | 0.03005% | 0.0289% | 0.1047% | 0%
19.2 | 28.34 | 29.85 | 31.6 | 32.99 | 34.4 | 38 | 39.5 | 19.2 | 28.28 | 29.93 | 31.63 | 33.04 | 34.41 | 38.04 | 39.47 | 19.2 | -0.2117% | 0.268% | 0.09494% | 0.1516% | 0.02907% | 0.1053% | -0.07595%
19.3 | 28.13 | 29.69 | 31.4 | 32.89 | 34.21 | 37.8 | 39.2 | 19.3 | 28.1 | 29.72 | 31.43 | 32.85 | 34.2 | 37.84 | 39.24 | 19.3 | -0.1066% | 0.101% | 0.09554% | -0.1216% | -0.02923% | 0.1058% | 0.102%
19.4 | 27.92 | 29.5 | 31.2 | 32.61 | 33.91 | 37.6 | 39 | 19.4 | 27.93 | 29.55 | 31.23 | 32.66 | 33.99 | 37.6 | 39.03 | 19.4 | 0.03582% | 0.1695% | 0.09615% | 0.1533% | 0.2359% | 0% | 0.07692%
19.5 | 27.76 | 29.4 | 30.98 | 32.41 | 33.78 | 37.3 | 38.8 | 19.5 | 27.76 | 29.37 | 31.01 | 32.47 | 33.77 | 37.37 | 38.81 | 19.5 | 0% | -0.102% | 0.09684% | 0.1851% | -0.0296% | 0.1877% | 0.02577%
19.6 | 27.58 | 29.1 | 30.82 | 32.21 | 33.51 | 37.1 | 38.6 | 19.6 | 27.56 | 29.16 | 30.8 | 32.25 | 33.52 | 37.19 | 38.61 | 19.6 | -0.07252% | 0.2062% | -0.06489% | 0.1242% | 0.02984% | 0.2426% | 0.02591%
19.7 | 27.31 | 28.9 | 30.49 | 32.01 | 33.3 | 36.92 | 38.4 | 19.7 | 27.35 | 28.95 | 30.57 | 32.05 | 33.32 | 37 | 38.41 | 19.7 | 0.1465% | 0.173% | 0.2624% | 0.125% | 0.06006% | 0.2167% | 0.02604%
19.8 | 27.1 | 28.7 | 30.41 | 31.81 | 33.2 | 36.7 | 38.2 | 19.8 | 27.12 | 28.75 | 30.4 | 31.84 | 33.13 | 36.75 | 38.2 | 19.8 | 0.0738% | 0.1742% | -0.03288% | 0.09431% | -0.2108% | 0.1362% | 0%
19.9 | 26.93 | 28.5 | 30.2 | 31.51 | 32.91 | 36.5 | 37.9 | 19.9 | 26.92 | 28.55 | 30.23 | 31.63 | 32.91 | 36.54 | 37.97 | 19.9 | -0.03713% | 0.1754% | 0.09934% | 0.3808% | 0% | 0.1096% | 0.1847%
20 | 26.76 | 28.3 | 30 | 31.41 | 32.7 | 36.2 | 37.7 | 20 | 26.77 | 28.36 | 30.01 | 31.44 | 32.71 | 36.35 | 37.73 | 20 | 0.03737% | 0.212% | 0.03333% | 0.09551% | 0.03058% | 0.4144% | 0.07958%
20.1 | 26.62 | 28.2 | 29.69 | 31.21 | 32.51 | 36.13 | 37.4 | 20.1 | 26.62 | 28.2 | 29.77 | 31.25 | 32.52 | 36.15 | 37.48 | 20.1 | 0% | 0% | 0.2695% | 0.1282% | 0.03076% | 0.05536% | 0.2139%
20.2 | 26.46 | 27.9 | 29.61 | 31.01 | 32.32 | 35.92 | 37.2 | 20.2 | 26.47 | 28 | 29.59 | 31.05 | 32.32 | 35.95 | 37.28 | 20.2 | 0.03779% | 0.3584% | -0.06754% | 0.129% | 0% | 0.08352% | 0.2151%
20.3 | 26.26 | 27.8 | 29.4 | 30.81 | 32.02 | 35.64 | 37.1 | 20.3 | 26.27 | 27.81 | 29.42 | 30.84 | 32.1 | 35.7 | 37.08 | 20.3 | 0.03808% | 0.03597% | 0.06803% | 0.09737% | 0.2498% | 0.1684% | -0.05391%
20.4 | 26.06 | 27.6 | 29.18 | 30.51 | 31.82 | 35.42 | 36.9 | 20.4 | 26.07 | 27.65 | 29.21 | 30.63 | 31.9 | 35.46 | 36.89 | 20.4 | 0.03837% | 0.1812% | 0.1028% | 0.3933% | 0.2514% | 0.1129% | -0.0271%
20.5 | 25.85 | 27.4 | 29.02 | 30.41 | 31.72 | 35.22 | 36.7 | 20.5 | 25.89 | 27.43 | 29.02 | 30.44 | 31.72 | 35.26 | 36.69 | 20.5 | 0.1547% | 0.1095% | 0% | 0.09865% | 0% | 0.1136% | -0.02725%
20.6 | 25.77 | 27.2 | 28.82 | 30.21 | 31.52 | 35.02 | 36.52 | 20.6 | 25.75 | 27.21 | 28.84 | 30.23 | 31.51 | 35.06 | 36.47 | 20.6 | -0.07761% | 0.03676% | 0.0694% | 0.0662% | -0.03173% | 0.1142% | -0.1369%
20.7 | 25.56 | 27 | 28.62 | 30.01 | 31.32 | 34.83 | 36.31 | 20.7 | 25.57 | 27.03 | 28.64 | 30.03 | 31.31 | 34.84 | 36.27 | 20.7 | 0.03912% | 0.1111% | 0.06988% | 0.06664% | -0.03193% | 0.02871% | -0.1102%
20.8 | 25.37 | 26.8 | 28.42 | 29.81 | 31.12 | 34.6 | 36.1 | 20.8 | 25.37 | 26.87 | 28.44 | 29.85 | 31.12 | 34.64 | 36.06 | 20.8 | 0% | 0.2612% | 0.07037% | 0.1342% | 0% | 0.1156% | -0.1108%
20.9 | 25.24 | 26.6 | 28.22 | 29.61 | 30.82 | 34.42 | 35.82 | 20.9 | 25.2 | 26.71 | 28.28 | 29.64 | 30.91 | 34.46 | 35.84 | 20.9 | -0.1585% | 0.4135% | 0.2126% | 0.1013% | 0.292% | 0.1162% | 0.05583%
21 | 25.01 | 26.5 | 28.12 | 29.41 | 30.72 | 34.22 | 35.62 | 21 | 25.04 | 26.54 | 28.13 | 29.44 | 30.71 | 34.26 | 35.63 | 21 | 0.12% | 0.1509% | 0.03556% | 0.102% | -0.03255% | 0.1169% | 0.02807%
21.1 | 24.86 | 26.3 | 27.82 | 29.21 | 30.52 | 34.02 | 35.42 | 21.1 | 24.87 | 26.37 | 27.93 | 29.24 | 30.51 | 34.06 | 35.43 | 21.1 | 0.04023% | 0.2662% | 0.3954% | 0.1027% | -0.03277% | 0.1176% | 0.02823%
21.2 | 24.72 | 26.1 | 27.72 | 29.01 | 30.32 | 33.82 | 35.32 | 21.2 | 24.72 | 26.21 | 27.74 | 29.08 | 30.35 | 33.86 | 35.26 | 21.2 | 0% | 0.4215% | 0.07215% | 0.2413% | 0.09894% | 0.1183% | -0.1699%
21.3 | 24.56 | 26 | 27.52 | 28.91 | 30.12 | 33.62 | 35.1 | 21.3 | 24.56 | 26.04 | 27.54 | 28.94 | 30.17 | 33.65 | 35.1 | 21.3 | 0% | 0.1538% | 0.07267% | 0.1038% | 0.166% | 0.08923% | 0%
21.4 | 24.36 | 25.8 | 27.32 | 28.61 | 29.92 | 33.41 | 34.83 | 21.4 | 24.37 | 25.85 | 27.35 | 28.75 | 29.95 | 33.49 | 34.91 | 21.4 | 0.04105% | 0.1938% | 0.1098% | 0.4893% | 0.1003% | 0.2394% | 0.2297%
21.5 | 24.23 | 25.62 | 27.22 | 28.54 | 29.72 | 33.33 | 34.72 | 21.5 | 24.22 | 25.66 | 27.18 | 28.55 | 29.76 | 33.31 | 34.71 | 21.5 | -0.04127% | 0.1561% | -0.147% | 0.03504% | 0.1346% | -0.06001% | -0.0288%
21.6 | 24.07 | 25.51 | 27.02 | 28.33 | 29.62 | 33.02 | 34.44 | 21.6 | 24.08 | 25.53 | 27.02 | 28.36 | 29.62 | 33.07 | 34.48 | 21.6 | 0.04155% | 0.0784% | 0% | 0.1059% | 0% | 0.1514% | 0.1161%
21.7 | 23.92 | 25.41 | 26.82 | 28.12 | 29.43 | 32.81 | 34.31 | 21.7 | 23.93 | 25.42 | 26.86 | 28.18 | 29.46 | 32.88 | 34.29 | 21.7 | 0.04181% | 0.03935% | 0.1491% | 0.2134% | 0.1019% | 0.2133% | -0.05829%
21.8 | 23.77 | 25.22 | 26.62 | 27.94 | 29.16 | 32.73 | 34.12 | 21.8 | 23.8 | 25.26 | 26.7 | 28.02 | 29.26 | 32.74 | 34.13 | 21.8 | 0.1262% | 0.1586% | 0.3005% | 0.2863% | 0.3429% | 0.03055% | 0.02931%
21.9 | 23.69 | 25.03 | 26.52 | 27.84 | 29.05 | 32.52 | 33.92 | 21.9 | 23.66 | 25.05 | 26.53 | 27.85 | 29.05 | 32.55 | 33.93 | 21.9 | -0.1266% | 0.0799% | 0.03771% | 0.03592% | 0% | 0.09225% | 0.02948%
22 | 23.43 | 24.8 | 26.32 | 27.63 | 28.94 | 32.23 | 33.72 | 22 | 23.47 | 24.85 | 26.34 | 27.66 | 28.88 | 32.35 | 33.73 | 22 | 0.1707% | 0.2016% | 0.07599% | 0.1086% | -0.2073% | 0.3723% | 0.02966%
22.1 | 23.3 | 24.72 | 26.12 | 27.43 | 28.72 | 32.15 | 33.52 | 22.1 | 23.27 | 24.7 | 26.14 | 27.46 | 28.72 | 32.16 | 33.53 | 22.1 | -0.1288% | -0.08091% | 0.07657% | 0.1094% | 0% | 0.0311% | 0.02983%
22.2 | 23.12 | 24.5 | 25.92 | 27.22 | 28.54 | 31.94 | 33.32 | 22.2 | 23.13 | 24.54 | 25.96 | 27.28 | 28.55 | 31.97 | 33.33 | 22.2 | 0.04325% | 0.1633% | 0.1543% | 0.2204% | 0.03504% | 0.09393% | 0.03001%
22.3 | 22.98 | 24.32 | 25.72 | 27.05 | 28.34 | 31.74 | 33.11 | 22.3 | 22.98 | 24.36 | 25.8 | 27.09 | 28.35 | 31.78 | 33.15 | 22.3 | 0% | 0.1645% | 0.311% | 0.1479% | 0.03529% | 0.126% | 0.1208%
22.4 | 22.82 | 24.11 | 25.64 | 26.92 | 28.13 | 31.61 | 32.94 | 22.4 | 22.83 | 24.19 | 25.64 | 26.92 | 28.18 | 31.59 | 32.95 | 22.4 | 0.04382% | 0.3318% | 0% | 0% | 0.1777% | -0.06327% | 0.03036%
22.5 | 22.66 | 24.03 | 25.43 | 26.73 | 28.04 | 31.32 | 32.75 | 22.5 | 22.69 | 24.06 | 25.49 | 26.77 | 28.03 | 31.41 | 32.75 | 22.5 | 0.1324% | 0.1248% | 0.2359% | 0.1496% | -0.03566% | 0.2874% | 0%
22.6 | 22.54 | 23.91 | 25.34 | 26.61 | 27.84 | 31.25 | 32.6 | 22.6 | 22.53 | 23.92 | 25.34 | 26.62 | 27.85 | 31.25 | 32.6 | 22.6 | -0.04437% | 0.04182% | 0% | 0.03758% | 0.03592% | 0% | 0%
22.7 | 22.39 | 23.72 | 25.14 | 26.44 | 27.63 | 31.04 | 32.44 | 22.7 | 22.37 | 23.76 | 25.16 | 26.44 | 27.68 | 31.07 | 32.45 | 22.7 | -0.08933% | 0.1686% | 0.07955% | 0% | 0.181% | 0.09665% | 0.03083%
22.8 | 22.22 | 23.46 | 25.02 | 26.21 | 27.55 | 30.84 | 32.31 | 22.8 | 22.23 | 23.56 | 25.02 | 26.25 | 27.5 | 30.87 | 32.27 | 22.8 | 0.045% | 0.4263% | 0% | 0.1526% | -0.1815% | 0.09728% | -0.1238%
22.9 | 22.09 | 23.41 | 24.84 | 26.12 | 27.23 | 30.64 | 32.02 | 22.9 | 22.09 | 23.42 | 24.86 | 26.12 | 27.29 | 30.67 | 32.08 | 22.9 | 0% | 0.04272% | 0.08052% | 0% | 0.2203% | 0.09791% | 0.1874%
23 | 21.91 | 23.23 | 24.63 | 25.93 | 27.14 | 30.43 | 31.95 | 23 | 21.94 | 23.3 | 24.69 | 25.96 | 27.13 | 30.5 | 31.93 | 23 | 0.1369% | 0.3013% | 0.2436% | 0.1157% | -0.03685% | 0.23% | -0.0626%
23.1 | 21.81 | 23.15 | 24.54 | 25.72 | 26.93 | 30.35 | 31.67 | 23.1 | 21.83 | 23.16 | 24.55 | 25.79 | 26.97 | 30.35 | 31.72 | 23.1 | 0.0917% | 0.0432% | 0.04075% | 0.2722% | 0.1485% | 0% | 0.1579%
23.2 | 21.68 | 22.95 | 24.42 | 25.63 | 26.75 | 30.14 | 31.48 | 23.2 | 21.69 | 22.99 | 24.42 | 25.64 | 26.8 | 30.17 | 31.52 | 23.2 | 0.04613% | 0.1743% | 0% | 0.03902% | 0.1869% | 0.09954% | 0.1271%
23.3 | 21.49 | 22.8 | 24.24 | 25.43 | 26.55 | 29.94 | 31.42 | 23.3 | 21.51 | 22.83 | 24.26 | 25.46 | 26.61 | 29.97 | 31.39 | 23.3 | 0.09307% | 0.1316% | 0.08251% | 0.118% | 0.226% | 0.1002% | -0.09548%
23.4 | 21.4 | 22.63 | 24.03 | 25.23 | 26.44 | 29.74 | 31.25 | 23.4 | 21.37 | 22.7 | 24.09 | 25.27 | 26.46 | 29.78 | 31.23 | 23.4 | -0.1402% | 0.3093% | 0.2497% | 0.1585% | 0.07564% | 0.1345% | -0.064%
23.5 | 21.22 | 22.55 | 23.94 | 25.11 | 26.34 | 29.6 | 31.01 | 23.5 | 21.24 | 22.56 | 23.94 | 25.12 | 26.33 | 29.62 | 31.01 | 23.5 | 0.09425% | 0.04435% | 0% | 0.03982% | -0.03797% | 0.06757% | 0%
23.6 | 21.1 | 22.33 | 23.74 | 24.92 | 26.16 | 29.44 | 30.79 | 23.6 | 21.1 | 22.4 | 23.76 | 24.99 | 26.15 | 29.47 | 30.82 | 23.6 | 0% | 0.3135% | 0.08425% | 0.2809% | -0.03823% | 0.1019% | 0.09743%
23.7 | 20.92 | 22.24 | 23.62 | 24.83 | 26.02 | 29.24 | 30.64 | 23.7 | 20.94 | 22.28 | 23.61 | 24.84 | 26.01 | 29.27 | 30.65 | 23.7 | 0.0956% | 0.1799% | -0.04234% | 0.04027% | -0.03843% | 0.1026% | 0.03264%
23.8 | 20.77 | 22.08 | 23.43 | 24.63 | 25.74 | 29.12 | 30.52 | 23.8 | 20.8 | 22.13 | 23.45 | 24.66 | 25.84 | 29.1 | 30.48 | 23.8 | 0.1444% | 0.2264% | 0.08536% | 0.1218% | 0.3885% | -0.06868% | -0.1311%
23.9 | 20.71 | 21.96 | 23.23 | 24.51 | 25.66 | 28.89 | 30.3 | 23.9 | 20.67 | 21.97 | 23.29 | 24.52 | 25.67 | 28.94 | 30.3 | 23.9 | -0.1931% | 0.04554% | 0.2583% | 0.0408% | 0.03897% | 0.1731% | 0%
24 | 20.52 | 21.8 | 23.14 | 24.34 | 25.57 | 28.74 | 30.07 | 24 | 20.54 | 21.83 | 23.14 | 24.37 | 25.54 | 28.77 | 30.12 | 24 | 0.09747% | 0.1376% | 0% | 0.1233% | -0.1173% | 0.1044% | 0.1663%
24.1 | 20.42 | 21.63 | 22.93 | 24.14 | 25.34 | 28.53 | 29.91 | 24.1 | 20.42 | 21.7 | 22.98 | 24.2 | 25.38 | 28.59 | 29.93 | 24.1 | 0% | 0.3236% | 0.2181% | 0.2486% | 0.1579% | 0.2103% | 0.06687%
24.2 | 20.2 | 21.55 | 22.84 | 24.06 | 25.19 | 28.45 | 29.74 | 24.2 | 20.27 | 21.56 | 22.84 | 24.06 | 25.22 | 28.45 | 29.75 | 24.2 | 0.3465% | 0.0464% | 0% | 0% | 0.1191% | 0% | 0.03362%
24.3 | 20.11 | 21.33 | 22.63 | 23.84 | 25.07 | 28.24 | 29.65 | 24.3 | 20.14 | 21.4 | 22.68 | 23.9 | 25.05 | 28.27 | 29.63 | 24.3 | 0.1492% | 0.3282% | 0.2209% | 0.2517% | -0.07978% | 0.1062% | -0.06745%
24.4 | 19.99 | 21.25 | 22.54 | 23.76 | 24.86 | 28.04 | 29.45 | 24.4 | 20.01 | 21.26 | 22.55 | 23.76 | 24.86 | 28.08 | 29.45 | 24.4 | 0.1001% | 0.04706% | 0.04437% | 0% | 0% | 0.1427% | 0%
24.5 | 19.91 | 21.1 | 22.42 | 23.54 | 24.65 | 27.9 | 29.32 | 24.5 | 19.87 | 21.13 | 22.41 | 23.6 | 24.69 | 27.93 | 29.28 | 24.5 | -0.2009% | 0.1422% | -0.0446% | 0.2549% | 0.1623% | 0.1075% | -0.1364%
24.6 | 19.71 | 20.94 | 22.23 | 23.46 | 24.55 | 27.73 | 29.09 | 24.6 | 19.74 | 20.98 | 22.29 | 23.46 | 24.57 | 27.8 | 29.12 | 24.6 | 0.1522% | 0.191% | 0.2699% | 0% | 0.08147% | 0.2524% | 0.1031%
24.7 | 19.59 | 20.8 | 22.16 | 23.31 | 24.47 | 27.67 | 28.92 | 24.7 | 19.61 | 20.83 | 22.15 | 23.33 | 24.44 | 27.66 | 28.96 | 24.7 | 0.1021% | 0.1442% | -0.04513% | 0.0858% | -0.1226% | -0.03614% | 0.1383%
24.8 | 19.51 | 20.64 | 21.95 | 23.15 | 24.25 | 27.46 | 28.78 | 24.8 | 19.47 | 20.68 | 22 | 23.18 | 24.29 | 27.49 | 28.8 | 24.8 | -0.205% | 0.1938% | 0.2278% | 0.1296% | 0.1649% | 0.1092% | 0.06949%
24.9 | 19.31 | 20.51 | 21.86 | 23.01 | 24.16 | 27.25 | 28.65 | 24.9 | 19.34 | 20.53 | 21.85 | 23.03 | 24.14 | 27.31 | 28.63 | 24.9 | 0.1554% | 0.09751% | -0.04575% | 0.08692% | -0.08278% | 0.2202% | -0.06981%
25 | 19.21 | 20.41 | 21.65 | 22.83 | 23.94 | 27.17 | 28.43 | 25 | 19.24 | 20.43 | 21.7 | 22.89 | 23.98 | 27.16 | 28.46 | 25 | 0.1562% | 0.09799% | 0.2309% | 0.2628% | 0.1671% | -0.03681% | 0.1055%
25.1 | 19.11 | 20.31 | 21.57 | 22.68 | 23.79 | 26.95 | 28.28 | 25.1 | 19.13 | 20.32 | 21.56 | 22.73 | 23.81 | 27.01 | 28.3 | 25.1 | 0.1047% | 0.04924% | -0.04636% | 0.2205% | 0.08407% | 0.2226% | 0.07072%
25.2 | 18.99 | 20.16 | 21.42 | 22.55 | 23.67 | 26.87 | 28.16 | 25.2 | 18.99 | 20.17 | 21.43 | 22.58 | 23.65 | 26.86 | 28.14 | 25.2 | 0% | 0.0496% | 0.04669% | 0.133% | -0.0845% | -0.03722% | -0.07102%
25.3 | 18.81 | 19.98 | 21.24 | 22.46 | 23.52 | 26.65 | 28.02 | 25.3 | 18.84 | 20.04 | 21.29 | 22.46 | 23.52 | 26.71 | 27.99 | 25.3 | 0.1595% | 0.3003% | 0.2354% | 0% | 0% | 0.2251% | -0.1071%
25.4 | 18.72 | 19.85 | 21.17 | 22.25 | 23.36 | 26.57 | 27.79 | 25.4 | 18.74 | 19.91 | 21.16 | 22.28 | 23.4 | 26.56 | 27.84 | 25.4 | 0.1068% | 0.3023% | -0.04724% | 0.1348% | 0.1712% | -0.03764% | 0.1799%
25.5 | 18.62 | 19.76 | 21.02 | 22.11 | 23.28 | 26.36 | 27.71 | 25.5 | 18.64 | 19.79 | 21.02 | 22.13 | 23.28 | 26.39 | 27.69 | 25.5 | 0.1074% | 0.1518% | 0% | 0.09046% | 0% | 0.1138% | -0.07218%
25.6 | 18.52 | 19.67 | 20.84 | 22.01 | 23.19 | 26.2 | 27.49 | 25.6 | 18.52 | 19.67 | 20.89 | 22.03 | 23.16 | 26.23 | 27.54 | 25.6 | 0% | 0% | 0.2399% | 0.09087% | -0.1294% | 0.1145% | 0.1819%
25.7 | 18.42 | 19.49 | 20.75 | 21.92 | 22.97 | 26.06 | 27.4 | 25.7 | 18.39 | 19.56 | 20.78 | 21.92 | 23 | 26.08 | 27.41 | 25.7 | -0.1629% | 0.3592% | 0.1446% | 0% | 0.1306% | 0.07675% | 0.0365%
25.8 | 18.32 | 19.46 | 20.65 | 21.69 | 22.89 | 25.85 | 27.26 | 25.8 | 18.29 | 19.48 | 20.67 | 21.75 | 22.86 | 25.91 | 27.26 | 25.8 | -0.1638% | 0.1028% | 0.09685% | 0.2766% | -0.1311% | 0.2321% | 0%
25.9 | 18.22 | 19.33 | 20.49 | 21.6 | 22.71 | 25.77 | 27.11 | 25.9 | 18.18 | 19.34 | 20.52 | 21.63 | 22.73 | 25.77 | 27.09 | 25.9 | -0.2195% | 0.05173% | 0.1464% | 0.1389% | 0.08807% | 0% | -0.07377%
26 | 18.02 | 19.18 | 20.36 | 21.45 | 22.56 | 25.6 | 26.84 | 26 | 18.05 | 19.18 | 20.38 | 21.5 | 22.59 | 25.64 | 26.9 | 26 | 0.1665% | 0% | 0.09823% | 0.2331% | 0.133% | 0.1562% | 0.2235%
26.1 | 17.92 | 19 | 20.27 | 21.33 | 22.44 | 25.46 | 26.77 | 26.1 | 17.92 | 19.04 | 20.26 | 21.35 | 22.43 | 25.49 | 26.74 | 26.1 | 0% | 0.2105% | -0.04933% | 0.09376% | -0.04456% | 0.1178% | -0.1121%
26.2 | 17.81 | 18.9 | 20.12 | 21.22 | 22.28 | 25.27 | 26.55 | 26.2 | 17.82 | 18.94 | 20.13 | 21.22 | 22.28 | 25.32 | 26.59 | 26.2 | 0.05615% | 0.2116% | 0.0497% | 0% | 0% | 0.1979% | 0.1507%
26.3 | 17.71 | 18.8 | 20.01 | 21.06 | 22.19 | 25.19 | 26.47 | 26.3 | 17.72 | 18.84 | 20.02 | 21.11 | 22.16 | 25.18 | 26.44 | 26.3 | 0.05647% | 0.2128% | 0.04998% | 0.2374% | -0.1352% | -0.0397% | -0.1133%
26.4 | 17.61 | 18.65 | 19.84 | 20.98 | 22.01 | 25 | 26.3 | 26.4 | 17.58 | 18.71 | 19.89 | 20.97 | 22.03 | 25.05 | 26.31 | 26.4 | -0.1704% | 0.3217% | 0.252% | -0.04766% | 0.09087% | 0.2% | 0.03802%
26.5 | 17.41 | 18.56 | 19.76 | 20.81 | 21.88 | 24.87 | 26.16 | 26.5 | 17.46 | 18.59 | 19.76 | 20.84 | 21.88 | 24.91 | 26.16 | 26.5 | 0.2872% | 0.1616% | 0% | 0.1442% | 0% | 0.1608% | 0%
26.6 | 17.41 | 18.47 | 19.61 | 20.66 | 21.72 | 24.79 | 26 | 26.6 | 17.37 | 18.47 | 19.63 | 20.71 | 21.73 | 24.78 | 26.01 | 26.6 | -0.2298% | 0% | 0.102% | 0.242% | 0.04604% | -0.04034% | 0.03846%
26.7 | 17.21 | 18.25 | 19.46 | 20.58 | 21.61 | 24.58 | 25.85 | 26.7 | 17.26 | 18.31 | 19.5 | 20.57 | 21.63 | 24.62 | 25.88 | 26.7 | 0.2905% | 0.3288% | 0.2055% | -0.04859% | 0.09255% | 0.1627% | 0.1161%
26.8 | 17.21 | 18.16 | 19.38 | 20.41 | 21.52 | 24.49 | 25.77 | 26.8 | 17.17 | 18.19 | 19.38 | 20.44 | 21.5 | 24.48 | 25.74 | 26.8 | -0.2324% | 0.1652% | 0% | 0.147% | -0.09294% | -0.04083% | -0.1164%
26.9 | 17.01 | 18.06 | 19.29 | 20.3 | 21.3 | 24.28 | 25.55 | 26.9 | 17.03 | 18.09 | 19.27 | 20.33 | 21.35 | 24.32 | 25.59 | 26.9 | 0.1176% | 0.1661% | -0.1037% | 0.1478% | 0.2347% | 0.1647% | 0.1566%
27 | 16.91 | 17.97 | 19.1 | 20.16 | 21.21 | 24.19 | 25.48 | 27 | 16.89 | 17.99 | 19.16 | 20.21 | 21.24 | 24.17 | 25.45 | 27 | -0.1183% | 0.1113% | 0.3141% | 0.248% | 0.1414% | -0.08268% | -0.1177%
27.1 | 16.81 | 17.88 | 19.08 | 20.08 | 21.14 | 23.98 | 25.26 | 27.1 | 16.82 | 17.88 | 19.06 | 20.07 | 21.12 | 24.02 | 25.27 | 27.1 | 0.05949% | 0% | -0.1048% | -0.0498% | -0.09461% | 0.1668% | 0.03959%
27.2 | 16.71 | 17.7 | 18.91 | 19.91 | 21.01 | 23.89 | 25.1 | 27.2 | 16.72 | 17.75 | 18.94 | 19.94 | 20.97 | 23.89 | 25.12 | 27.2 | 0.05984% | 0.2825% | 0.1586% | 0.1507% | -0.1904% | 0% | 0.07968%
27.3 | 16.61 | 17.6 | 18.81 | 19.8 | 20.81 | 23.81 | 24.99 | 27.3 | 16.59 | 17.65 | 18.83 | 19.83 | 20.84 | 23.8 | 25.02 | 27.3 | -0.1204% | 0.2841% | 0.1063% | 0.1515% | 0.1442% | -0.042% | 0.12%
27.4 | 16.52 | 17.47 | 18.67 | 19.66 | 20.71 | 23.59 | 24.9 | 27.4 | 16.5 | 17.52 | 18.69 | 19.7 | 20.73 | 23.67 | 24.92 | 27.4 | -0.1211% | 0.2862% | 0.1071% | 0.2035% | 0.09657% | 0.3391% | 0.08032%
27.5 | 16.42 | 17.39 | 18.51 | 19.57 | 20.59 | 23.48 | 24.78 | 27.5 | 16.4 | 17.4 | 18.53 | 19.58 | 20.61 | 23.5 | 24.78 | 27.5 | -0.1218% | 0.0575% | 0.108% | 0.0511% | 0.09713% | 0.08518% | 0%
27.6 | 16.32 | 17.29 | 18.36 | 19.48 | 20.5 | 23.3 | 24.58 | 27.6 | 16.3 | 17.3 | 18.4 | 19.47 | 20.47 | 23.35 | 24.6 | 27.6 | -0.1225% | 0.05784% | 0.2179% | -0.05133% | -0.1463% | 0.2146% | 0.08137%
27.7 | 16.22 | 17.19 | 18.27 | 19.31 | 20.31 | 23.17 | 24.49 | 27.7 | 16.19 | 17.2 | 18.28 | 19.34 | 20.34 | 23.21 | 24.47 | 27.7 | -0.185% | 0.05817% | 0.05473% | 0.1554% | 0.1477% | 0.1726% | -0.08167%
27.8 | 16.03 | 17.09 | 18.17 | 19.2 | 20.21 | 23.08 | 24.38 | 27.8 | 16.07 | 17.09 | 18.18 | 19.24 | 20.23 | 23.09 | 24.37 | 27.8 | 0.2495% | 0% | 0.05504% | 0.2083% | 0.09896% | 0.04333% | -0.04102%
27.9 | 16.02 | 16.99 | 18.08 | 19.06 | 20.11 | 22.99 | 24.27 | 27.9 | 15.98 | 16.99 | 18.08 | 19.11 | 20.13 | 22.97 | 24.23 | 27.9 | -0.2497% | 0% | 0% | 0.2623% | 0.09945% | -0.08699% | -0.1648%
28 | 15.81 | 16.89 | 17.98 | 18.97 | 19.99 | 22.78 | 24.09 | 28 | 15.86 | 16.89 | 17.98 | 18.98 | 20.01 | 22.82 | 24.08 | 28 | 0.3163% | 0% | 0% | 0.05271% | 0.1001% | 0.1756% | -0.04151%
28.1 | 15.71 | 16.79 | 17.89 | 18.9 | 19.9 | 22.69 | 23.99 | 28.1 | 15.75 | 16.78 | 17.87 | 18.88 | 19.87 | 22.68 | 23.96 | 28.1 | 0.2546% | -0.05956% | -0.1118% | -0.1058% | -0.1508% | -0.04407% | -0.1251%
28.2 | 15.71 | 16.6 | 17.71 | 18.71 | 19.71 | 22.5 | 23.78 | 28.2 | 15.68 | 16.65 | 17.74 | 18.75 | 19.74 | 22.55 | 23.8 | 28.2 | -0.191% | 0.3012% | 0.1694% | 0.2138% | 0.1522% | 0.2222% | 0.0841%
28.3 | 15.61 | 16.5 | 17.61 | 18.6 | 19.59 | 22.38 | 23.69 | 28.3 | 15.59 | 16.54 | 17.64 | 18.64 | 19.61 | 22.41 | 23.66 | 28.3 | -0.1281% | 0.2424% | 0.1704% | 0.2151% | 0.1021% | 0.134% | -0.1266%
28.4 | 15.51 | 16.37 | 17.51 | 18.5 | 19.52 | 22.3 | 23.5 | 28.4 | 15.49 | 16.42 | 17.54 | 18.54 | 19.51 | 22.33 | 23.53 | 28.4 | -0.1289% | 0.3054% | 0.1713% | 0.2162% | -0.05123% | 0.1345% | 0.1277%
28.5 | 15.41 | 16.29 | 17.41 | 18.4 | 19.39 | 22.1 | 23.4 | 28.5 | 15.39 | 16.29 | 17.43 | 18.44 | 19.39 | 22.25 | 23.42 | 28.5 | -0.1298% | 0% | 0.1149% | 0.2174% | 0% | 0.6787% | 0.08547%
28.6 | 15.31 | 16.2 | 17.26 | 18.29 | 19.21 | 22.08 | 23.27 | 28.6 | 15.29 | 16.23 | 17.3 | 18.3 | 19.25 | 22.11 | 23.29 | 28.6 | -0.1306% | 0.1852% | 0.2317% | 0.05467% | 0.2082% | 0.1359% | 0.08595%
28.7 | 15.21 | 16.19 | 17.2 | 18.1 | 19.2 | 21.9 | 23.19 | 28.7 | 15.19 | 16.17 | 17.19 | 18.16 | 19.16 | 21.93 | 23.16 | 28.7 | -0.1315% | -0.1235% | -0.05814% | 0.3315% | -0.2083% | 0.137% | -0.1294%
28.8 | 15.11 | 15.99 | 17.1 | 18.09 | 19.01 | 21.8 | 23 | 28.8 | 15.09 | 16.06 | 17.09 | 18.07 | 19.04 | 21.8 | 23.03 | 28.8 | -0.1324% | 0.4378% | -0.05848% | -0.1106% | 0.1578% | 0% | 0.1304%
28.9 | 15.01 | 15.99 | 17 | 17.9 | 18.88 | 21.7 | 22.87 | 28.9 | 14.98 | 15.97 | 16.99 | 17.95 | 18.9 | 21.7 | 22.89 | 28.9 | -0.1999% | -0.1251% | -0.05882% | 0.2793% | 0.1059% | 0% | 0.08745%
29 | 14.81 | 15.8 | 16.9 | 17.8 | 18.89 | 21.6 | 22.79 | 29 | 14.86 | 15.85 | 16.89 | 17.84 | 18.77 | 21.59 | 22.76 | 29 | 0.3376% | 0.3165% | -0.05917% | 0.2247% | -0.6353% | -0.0463% | -0.1316%
29.1 | 14.81 | 15.7 | 16.8 | 17.7 | 18.7 | 21.4 | 22.6 | 29.1 | 14.81 | 15.74 | 16.79 | 17.76 | 18.68 | 21.46 | 22.63 | 29.1 | 0% | 0.2548% | -0.05952% | 0.339% | -0.107% | 0.2804% | 0.1327%
29.2 | 14.71 | 15.6 | 16.7 | 17.69 | 18.6 | 21.3 | 22.5 | 29.2 | 14.74 | 15.64 | 16.69 | 17.67 | 18.57 | 21.32 | 22.52 | 29.2 | 0.2039% | 0.2564% | -0.05988% | -0.1131% | -0.1613% | 0.0939% | 0.08889%
29.3 | 14.61 | 15.49 | 16.6 | 17.48 | 18.41 | 21.2 | 22.38 | 29.3 | 14.62 | 15.56 | 16.59 | 17.52 | 18.44 | 21.19 | 22.38 | 29.3 | 0.06845% | 0.4519% | -0.06024% | 0.2288% | 0.163% | -0.04717% | 0%
29.4 | 14.51 | 15.49 | 16.5 | 17.41 | 18.31 | 21 | 22.2 | 29.4 | 14.52 | 15.47 | 16.48 | 17.42 | 18.33 | 21.06 | 22.23 | 29.4 | 0.06892% | -0.1291% | -0.1212% | 0.05744% | 0.1092% | 0.2857% | 0.1351%
29.5 | 14.41 | 15.29 | 16.3 | 17.28 | 18.21 | 20.9 | 22.1 | 29.5 | 14.42 | 15.36 | 16.36 | 17.32 | 18.23 | 20.95 | 22.12 | 29.5 | 0.0694% | 0.4578% | 0.3681% | 0.2315% | 0.1098% | 0.2392% | 0.0905%
29.6 | 14.31 | 15.29 | 16.29 | 17.21 | 18.11 | 20.8 | 22 | 29.6 | 14.32 | 15.29 | 16.28 | 17.22 | 18.14 | 20.82 | 22.02 | 29.6 | 0.06988% | 0% | -0.06139% | 0.05811% | 0.1657% | 0.09615% | 0.09091%
29.7 | 14.21 | 15.2 | 16.2 | 17.08 | 18 | 20.7 | 21.89 | 29.7 | 14.25 | 15.19 | 16.19 | 17.11 | 18.04 | 20.7 | 21.92 | 29.7 | 0.2815% | -0.06579% | -0.06173% | 0.1756% | 0.2222% | 0% | 0.137%
29.8 | 14.21 | 14.99 | 16.1 | 17.01 | 17.91 | 20.6 | 21.77 | 29.8 | 14.18 | 15.06 | 16.09 | 17.02 | 17.91 | 20.6 | 21.79 | 29.8 | -0.2111% | 0.467% | -0.06211% | 0.05879% | 0% | 0% | 0.09187%
29.9 | 14.11 | 14.9 | 16 | 16.88 | 17.81 | 20.5 | 21.69 | 29.9 | 14.09 | 14.96 | 15.99 | 16.91 | 17.79 | 20.49 | 21.66 | 29.9 | -0.1417% | 0.4027% | -0.0625% | 0.1777% | -0.1123% | -0.04878% | -0.1383%
30 | 14.01 | 14.9 | 15.9 | 16.82 | 17.71 | 20.4 | 21.5 | 30 | 13.99 | 14.89 | 15.89 | 16.82 | 17.69 | 20.36 | 21.53 | 30 | -0.1428% | -0.06711% | -0.06289% | 0% | -0.1129% | -0.1961% | 0.1395%

## Open Field Profiles at 1.3cm
Lateral Profiles
Open field
SSD = 90 cm
Depth = 1.3 cm
Scanning direction: Crossline (i.e. along the leaf travel direction)
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 2x2 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 28x28 cm2
-25.2
-25.1
-25
-24.9
-24.8
-24.7
-24.6
-24.5
-24.4
-24.3
-24.2
-24.1
-24
-23.9
-23.8
-23.7
-23.6
-23.5
-23.4
-23.3
-23.2
-23.1
-23
-22.9
-22.8
-22.7
-22.6
-22.5
-22.4
-22.3
-22.2
-22.1
-22
-21.9
-21.8
-21.7
-21.6
-21.5
-21.4
-21.3
-21.2
-21.1
-21
-20.9
-20.8
-20.7
-20.6
-20.5
-20.4
-20.3
-20.2
-20.1
-20
-19.9
-19.8
-19.7 | [6 empty] | 1.6
-19.6 | [6 empty] | 1.62
-19.5 | [6 empty] | 1.7
-19.4 | [6 empty] | 1.7
-19.3 | [6 empty] | 1.7
-19.2 | [6 empty] | 1.72
-19.1 | [6 empty] | 1.8
-19 | [6 empty] | 1.8
-18.9 | [6 empty] | 1.8
-18.8 | [6 empty] | 1.78
-18.7 | [6 empty] | 1.95
-18.6 | [6 empty] | 1.9
-18.5 | [6 empty] | 1.9
-18.4 | [6 empty] | 1.9
-18.3 | [6 empty] | 1.92
-18.2 | [6 empty] | 2
-18.1 | [6 empty] | 2
-18 | [6 empty] | 2.02
-17.9 | [6 empty] | 2.1
-17.8 | [6 empty] | 2.1
-17.7 | [6 empty] | 2.08
-17.6 | [6 empty] | 2.25
-17.5 | [6 empty] | 2.2
-17.4 | [6 empty] | 2.2
-17.3 | [6 empty] | 2.21
-17.2 | [6 empty] | 2.3
-17.1 | [6 empty] | 2.32
-17 | [6 empty] | 2.4
-16.9 | [6 empty] | 2.4
-16.8 | [6 empty] | 2.41
-16.7 | [6 empty] | 2.5
-16.6 | [6 empty] | 2.46
-16.5 | [6 empty] | 2.71
-16.4 | [6 empty] | 2.6
-16.3 | [6 empty] | 2.61
-16.2 | [6 empty] | 2.7
-16.1 | [6 empty] | 2.71
-16 | [5 empty] | 1.5 | 2.8
-15.9 | [5 empty] | 1.5 | 2.81
-15.8 | [5 empty] | 1.5 | 2.9
-15.7 | [5 empty] | 1.59 | 2.91
-15.6 | [5 empty] | 1.6 | 3.01
-15.5 | [5 empty] | 1.6 | 3.1
-15.4 | [5 empty] | 1.6 | 3.11
-15.3 | [5 empty] | 1.69 | 3.16
-15.2 | [5 empty] | 1.7 | 3.42
-15.1 | [5 empty] | 1.7 | 3.32
-15 | [5 empty] | 1.7 | 3.36
-14.9 | [5 empty] | 1.79 | 3.58
-14.8 | [5 empty] | 1.8 | 3.72
-14.7 | [5 empty] | 1.8 | 3.62
-14.6 | [5 empty] | 1.8 | 3.71
-14.5 | [5 empty] | 1.89 | 3.81
-14.4 | [5 empty] | 1.9 | 3.87
-14.3 | [5 empty] | 1.9 | 4.23
-14.2 | [5 empty] | 1.88 | 4.22
-14.1 | [5 empty] | 2.01 | 4.33
-14 | [5 empty] | 2 | 4.53
-13.9 | [5 empty] | 2 | 4.73
-13.8 | [5 empty] | 2.09 | 4.89
-13.7 | [5 empty] | 2.1 | 5.36
-13.6 | [5 empty] | 2.1 | 5.57
-13.5 | [5 empty] | 2.18 | 6.04
-13.4 | [5 empty] | 2.2 | 6.86
-13.3 | [5 empty] | 2.2 | 8.15
-13.2 | [5 empty] | 2.28 | 10.31
-13.1 | [5 empty] | 2.3 | 14.74
-13 | [5 empty] | 2.38 | 20.89
-12.9 | [5 empty] | 2.4 | 28.59
-12.8 | [5 empty] | 2.4 | 36.58
-12.7 | [5 empty] | 2.48 | 44.69
-12.6 | [5 empty] | 2.5 | 51.15
-12.5 | [5 empty] | 2.58 | 57.21
-12.4 | [5 empty] | 2.6 | 62.11
-12.3 | [5 empty] | 2.68 | 64.06
-12.2 | [5 empty] | 2.7 | 65.4
-12.1 | [5 empty] | 2.78 | 66.26
-12 | [5 empty] | 2.8 | 66.95
-11.9 | [5 empty] | 2.87 | 67.45
-11.8 | [5 empty] | 2.87 | 67.94
-11.7 | [5 empty] | 3.02 | 68.34
-11.6 | [5 empty] | 3.08 | 68.74
-11.5 |  |  |  |  | 0.8 | 3.1 | 69.15
-11.4 |  |  |  |  | 0.8 | 3.18 | 69.41
-11.3 |  |  |  |  | 0.8 | 3.27 | 69.85
-11.2 |  |  |  |  | 0.8 | 3.27 | 70.11
-11.1 |  |  |  |  | 0.8 | 3.38 | 70.55
-11 |  |  |  |  | 0.8 | 3.49 | 70.82
-10.9 |  |  |  |  | 0.8 | 3.59 | 71.11
-10.8 |  |  |  |  | 0.8 | 3.69 | 71.53
-10.7 |  |  |  |  | 0.88 | 3.9 | 71.85
-10.6 |  |  |  | 0.6 | 0.9 | 3.98 | 72.11
-10.5 |  |  |  | 0.6 | 0.9 | 4.05 | 72.55
-10.4 |  |  |  | 0.6 | 0.9 | 4.3 | 72.82
-10.3 |  |  |  | 0.6 | 0.9 | 4.46 | 73.12
-10.2 |  |  |  | 0.6 | 0.85 | 4.62 | 73.41
-10.1 |  |  |  | 0.6 | 0.89 | 4.94 | 73.85
-10 |  |  |  | 0.6 | 1.02 | 5.24 | 74.12
-9.9 |  |  |  | 0.59 | 1 | 5.59 | 74.42
-9.8 |  |  |  | 0.71 | 1 | 6.23 | 74.71
-9.7 |  |  | 0.4 | 0.7 | 1 | 7.07 | 75.14
-9.6 |  |  | 0.4 | 0.7 | 1 | 8.56 | 75.54
-9.5 |  |  | 0.4 | 0.7 | 1.08 | 11.35 | 75.84
-9.4 |  |  | 0.4 | 0.7 | 1.1 | 16.43 | 76.14
-9.3 |  |  | 0.4 | 0.7 | 1.1 | 23.63 | 76.45
-9.2 |  |  | 0.4 | 0.7 | 1.1 | 32.03 | 76.7
-9.1 |  |  | 0.39 | 0.69 | 1.18 | 41.91 | 77.04
-9 |  |  | 0.46 | 0.8 | 1.2 | 50.97 | 77.3
-8.9 |  |  | 0.51 | 0.8 | 1.2 | 60.06 | 77.7
-8.8 |  | 0.2 | 0.5 | 0.8 | 1.2 | 66.87 | 78.04
-8.7 |  | 0.2 | 0.5 | 0.8 | 1.15 | 71.61 | 78.3
-8.6 |  | 0.2 | 0.5 | 0.8 | 1.32 | 74.41 | 78.7
-8.5 |  | 0.2 | 0.5 | 0.8 | 1.3 | 76 | 78.9
-8.4 |  | 0.2 | 0.5 | 0.8 | 1.3 | 77.01 | 79.3
-8.3 |  | 0.2 | 0.5 | 0.88 | 1.38 | 77.76 | 79.6
-8.2 |  | 0.2 | 0.5 | 0.9 | 1.4 | 78.41 | 79.9
-8.1 |  | 0.2 | 0.5 | 0.9 | 1.4 | 78.9 | 80.3
-8 |  | 0.2 | 0.5 | 0.9 | 1.35 | 79.42 | 80.64
-7.9 |  | 0.2 | 0.49 | 0.9 | 1.52 | 79.79 | 80.9
-7.8 |  | 0.2 | 0.56 | 0.98 | 1.5 | 80.25 | 81.2
-7.7 |  | 0.2 | 0.61 | 1 | 1.58 | 80.59 | 81.6
-7.6 |  | 0.2 | 0.6 | 1 | 1.6 | 81.04 | 81.9
-7.5 |  | 0.2 | 0.6 | 1 | 1.6 | 81.32 | 82.2
-7.4 |  | 0.25 | 0.6 | 0.99 | 1.67 | 81.73 | 82.5
-7.3 |  | 0.31 | 0.6 | 1.11 | 1.7 | 82.13 | 82.8
-7.2 |  | 0.3 | 0.6 | 1.1 | 1.77 | 82.55 | 83.2
-7.1 |  | 0.3 | 0.59 | 1.1 | 1.8 | 82.84 | 83.5
-7 |  | 0.3 | 0.66 | 1.09 | 1.87 | 83.2 | 83.8
-6.9 |  | 0.3 | 0.71 | 1.21 | 1.9 | 83.52 | 84.2
-6.8 |  | 0.3 | 0.7 | 1.2 | 1.97 | 83.95 | 84.5
-6.7 |  | 0.3 | 0.7 | 1.2 | 2 | 84.22 | 84.8
-6.6 |  | 0.3 | 0.7 | 1.27 | 2.07 | 84.63 | 85.1
-6.5 |  | 0.3 | 0.7 | 1.3 | 2.17 | 85.05 | 85.5
-6.4 |  | 0.3 | 0.69 | 1.28 | 2.27 | 85.34 | 85.8
-6.3 |  | 0.3 | 0.76 | 1.41 | 2.3 | 85.62 | 86.1
-6.2 |  | 0.3 | 0.81 | 1.4 | 2.25 | 86.05 | 86.4
-6.1 |  | 0.29 | 0.8 | 1.47 | 2.64 | 86.32 | 86.7
-6 |  | 0.4 | 0.79 | 1.5 | 2.68 | 86.75 | 87
-5.9 |  | 0.4 | 0.86 | 1.57 | 2.77 | 87.02 | 87.4
-5.8 |  | 0.4 | 0.91 | 1.58 | 2.95 | 87.47 | 87.7
-5.7 |  | 0.4 | 0.9 | 1.71 | 3.14 | 87.69 | 88
-5.6 |  | 0.4 | 0.9 | 1.77 | 3.22 | 88.03 | 88.3
-5.5 |  | 0.4 | 0.96 | 1.78 | 3.79 | 88.34 | 88.7
-5.4 |  | 0.4 | 1 | 1.91 | 4.03 | 88.71 | 89
-5.3 |  | 0.4 | 1 | 1.97 | 4.57 | 89.04 | 89.4
-5.2 |  | 0.4 | 1.05 | 2.07 | 5.52 | 89.41 | 89.6
-5.1 |  | 0.43 | 1.11 | 2.15 | 6.94 | 89.73 | 89.9
-5 |  | 0.5 | 1.1 | 2.29 | 9.46 | 90.04 | 90.2
-4.9 | 0.1 | 0.5 | 1.16 | 2.47 | 14.69 | 90.34 | 90.5
-4.8 | 0.1 | 0.5 | 1.19 | 2.66 | 22.27 | 90.64 | 90.9
-4.7 | 0.1 | 0.5 | 1.27 | 2.87 | 32.59 | 90.94 | 91.2
-4.6 | 0.1 | 0.5 | 1.34 | 3.12 | 43.94 | 91.31 | 91.5
-4.5 | 0.1 | 0.6 | 1.41 | 3.6 | 55.51 | 91.63 | 91.8
-4.4 | 0.1 | 0.6 | 1.44 | 4.16 | 66.44 | 91.94 | 92.1
-4.3 | 0.1 | 0.6 | 1.58 | 4.96 | 76.14 | 92.24 | 92.4
-4.2 | 0.1 | 0.6 | 1.66 | 6.42 | 82.65 | 92.54 | 92.6
-4.1 | 0.2 | 0.6 | 1.76 | 9.02 | 86.81 | 92.84 | 92.9
-4 | 0.2 | 0.7 | 1.88 | 14.06 | 89.07 | 93.13 | 93.2
-3.9 | 0.2 | 0.7 | 2.04 | 21.91 | 90.62 | 93.36 | 93.5
-3.8 | 0.2 | 0.7 | 2.27 | 32.36 | 91.45 | 93.62 | 93.8
-3.7 | 0.2 | 0.8 | 2.41 | 43.92 | 92.18 | 94.04 | 94.1
-3.6 | 0.2 | 0.8 | 2.99 | 56.12 | 92.75 | 94.27 | 94.3
-3.5 | 0.2 | 0.9 | 3.29 | 67.55 | 93.3 | 94.54 | 94.6
-3.4 | 0.2 | 0.9 | 4.14 | 77.66 | 93.61 | 94.83 | 94.9
-3.3 | 0.2 | 1 | 5.26 | 84.47 | 94.08 | 95.05 | 95.1
-3.2 | 0.2 | 1.1 | 7.49 | 88.89 | 94.49 | 95.34 | 95.37
-3.1 | 0.19 | 1.1 | 11.79 | 91.38 | 94.94 | 95.61 | 95.66
-3 | 0.31 | 1.2 | 18.9 | 92.87 | 95.15 | 95.91 | 95.9
-2.9 | 0.3 | 1.3 | 28.74 | 93.86 | 95.56 | 96.15 | 96.17
-2.8 | 0.3 | 1.48 | 40.6 | 94.54 | 95.75 | 96.37 | 96.4
-2.7 | 0.3 | 1.67 | 53.23 | 95.19 | 96.16 | 96.59 | 96.6
-2.6 | 0.3 | 1.85 | 65.23 | 95.62 | 96.25 | 96.85 | 96.8
-2.5 | 0.36 | 2.24 | 76.09 | 96.02 | 96.75 | 97.06 | 97
-2.4 | 0.4 | 2.68 | 84.59 | 96.36 | 96.76 | 97.34 | 97.2
-2.3 | 0.39 | 3.4 | 89.96 | 96.73 | 97.31 | 97.43 | 97.45
-2.2 | 0.51 | 4.63 | 93.14 | 97.02 | 97.42 | 97.74 | 97.65
-2.1 | 0.49 | 7.53 | 94.76 | 97.25 | 97.61 | 97.83 | 97.9
-2 | 0.6 | 15.8 | 95.77 | 97.57 | 97.7 | 98.09 | 98
-1.9 | 0.71 | 33.32 | 96.51 | 97.75 | 98.22 | 98.22 | 98.25
-1.8 | 0.85 | 58.15 | 97.11 | 98.07 | 98.32 | 98.42 | 98.39
-1.7 | 0.92 | 77.82 | 97.57 | 98.27 | 98.46 | 98.58 | 98.6
-1.6 | 1.26 | 89.06 | 97.82 | 98.35 | 98.61 | 98.79 | 98.7
-1.5 | 1.59 | 93.97 | 98.21 | 98.56 | 98.64 | 98.99 | 98.8
-1.4 | 2.3 | 96.38 | 98.45 | 98.76 | 99.06 | 99.03 | 99
-1.3 | 3.37 | 97.2 | 98.63 | 98.96 | 99.06 | 99.28 | 99.1
-1.2 | 6.03 | 97.77 | 98.9 | 99.17 | 99.2 | 99.36 | 99.25
-1.1 | 12.9 | 98.3 | 99.09 | 99.25 | 99.36 | 99.4 | 99.4
-1 | 30.36 | 98.76 | 99.25 | 99.46 | 99.46 | 99.43 | 99.45
-0.9 | 54.88 | 98.94 | 99.33 | 99.49 | 99.55 | 99.61 | 99.55
-0.8 | 77.09 | 99.29 | 99.5 | 99.59 | 99.65 | 99.77 | 99.6
-0.7 | 89.8 | 99.53 | 99.6 | 99.7 | 99.75 | 99.78 | 99.7
-0.6 | 95.85 | 99.44 | 99.61 | 99.8 | 99.75 | 99.92 | 99.8
-0.5 | 98.14 | 99.65 | 99.81 | 99.92 | 99.95 | 99.9 | 99.85
-0.4 | 99.29 | 99.77 | 99.85 | 99.95 | 99.9 | 99.95 | 99.8
-0.3 | 99.62 | 99.88 | 99.82 | 100 | 99.95 | 100 | 99.95
-0.2 | 99.86 | 100 | 99.95 | 100 | 100 | 100 | 99.89
-0.1 | 99.86 | 99.96 | 100 | 100 | 100 | 100 | 100
0 | 100 | 99.94 | 99.99 | 100 | 100 | 100 | 100
0.1 | 99.86 | 99.94 | 100 | 100 | 100 | 100 | 100
0.2 | 99.86 | 99.98 | 99.95 | 100 | 100 | 100 | 99.89
0.3 | 99.62 | 99.83 | 99.82 | 100 | 99.95 | 100 | 99.95
0.4 | 99.29 | 99.78 | 99.85 | 99.95 | 99.9 | 99.95 | 99.8
0.5 | 98.14 | 99.65 | 99.81 | 99.92 | 99.95 | 99.9 | 99.85
0.6 | 95.85 | 99.44 | 99.61 | 99.8 | 99.75 | 99.92 | 99.8
0.7 | 89.8 | 99.53 | 99.6 | 99.7 | 99.75 | 99.78 | 99.7
0.8 | 77.09 | 99.29 | 99.5 | 99.6 | 99.65 | 99.77 | 99.6
0.9 | 54.88 | 98.94 | 99.33 | 99.48 | 99.55 | 99.6 | 99.55
1 | 30.36 | 98.76 | 99.25 | 99.46 | 99.46 | 99.44 | 99.45
1.1 | 12.9 | 98.3 | 99.09 | 99.25 | 99.36 | 99.4 | 99.4
1.2 | 6.03 | 97.78 | 98.9 | 99.16 | 99.2 | 99.36 | 99.25
1.3 | 3.39 | 97.2 | 98.63 | 98.96 | 99.06 | 99.28 | 99.1
1.4 | 2.25 | 96.38 | 98.45 | 98.76 | 99.06 | 99.03 | 99
1.5 | 1.57 | 93.97 | 98.21 | 98.56 | 98.64 | 98.99 | 98.8
1.6 | 1.24 | 89.06 | 97.82 | 98.35 | 98.61 | 98.79 | 98.7
1.7 | 0.92 | 77.82 | 97.57 | 98.27 | 98.46 | 98.58 | 98.6
1.8 | 0.85 | 58.15 | 97.11 | 98.07 | 98.32 | 98.42 | 98.4
1.9 | 0.71 | 33.32 | 96.51 | 97.75 | 98.22 | 98.22 | 98.25
2 | 0.6 | 15.8 | 95.77 | 97.57 | 97.7 | 98.09 | 98
2.1 | 0.49 | 7.53 | 94.76 | 97.25 | 97.61 | 97.83 | 97.9
2.2 | 0.51 | 4.63 | 93.14 | 97.02 | 97.42 | 97.74 | 97.65
2.3 | 0.39 | 3.4 | 89.96 | 96.73 | 97.31 | 97.43 | 97.45
2.4 | 0.4 | 2.68 | 84.59 | 96.36 | 96.76 | 97.34 | 97.2
2.5 | 0.36 | 2.24 | 76.09 | 96.02 | 96.75 | 97.06 | 97
2.6 | 0.3 | 1.85 | 65.23 | 95.62 | 96.25 | 96.85 | 96.8
2.7 | 0.3 | 1.67 | 53.23 | 95.19 | 96.21 | 96.59 | 96.6
2.8 | 0.3 | 1.47 | 40.6 | 94.54 | 95.66 | 96.37 | 96.4
2.9 | 0.3 | 1.3 | 28.74 | 93.86 | 95.61 | 96.15 | 96.17
3 | 0.31 | 1.2 | 18.9 | 92.87 | 95.05 | 95.91 | 95.9
3.1 | 0.19 | 1.1 | 11.79 | 91.38 | 94.97 | 95.61 | 95.66
3.2 | 0.2 | 1.1 | 7.49 | 88.89 | 94.4 | 95.34 | 95.37
3.3 | 0.2 | 1 | 5.27 | 84.47 | 94.08 | 95.04 | 95.1
3.4 | 0.2 | 0.9 | 4.1 | 77.66 | 93.61 | 94.84 | 94.9
3.5 | 0.2 | 0.9 | 3.34 | 67.55 | 93.3 | 94.54 | 94.6
3.6 | 0.2 | 0.8 | 2.98 | 56.12 | 92.75 | 94.27 | 94.3
3.7 | 0.2 | 0.8 | 2.41 | 43.92 | 92.18 | 94.03 | 94.1
3.8 | 0.2 | 0.7 | 2.27 | 32.36 | 91.45 | 93.63 | 93.8
3.9 | 0.2 | 0.7 | 2.04 | 21.91 | 90.62 | 93.36 | 93.5
4 | 0.2 | 0.7 | 1.88 | 14.06 | 89.07 | 93.14 | 93.2
4.1 | 0.18 | 0.6 | 1.77 | 9.02 | 86.81 | 92.84 | 92.9
4.2 | 0.12 | 0.6 | 1.63 | 6.42 | 82.65 | 92.54 | 92.6
4.3 | 0.08 | 0.6 | 1.58 | 4.96 | 76.14 | 92.24 | 92.4
4.4 | 0.1 | 0.6 | 1.44 | 4.16 | 66.44 | 91.94 | 92.1
4.5 | 0.1 | 0.6 | 1.41 | 3.6 | 55.51 | 91.63 | 91.8
4.6 | 0.1 | 0.5 | 1.34 | 3.12 | 43.94 | 91.31 | 91.5
4.7 | 0.1 | 0.5 | 1.27 | 2.87 | 32.59 | 90.94 | 91.2
4.8 | 0.1 | 0.5 | 1.19 | 2.66 | 22.27 | 90.64 | 90.9
4.9 |  | 0.5 | 1.16 | 2.47 | 14.69 | 90.34 | 90.5
5 |  | 0.5 | 1.1 | 2.29 | 9.46 | 90.04 | 90.2
5.1 |  | 0.4 | 1.11 | 2.15 | 6.94 | 89.73 | 89.9
5.2 |  | 0.4 | 1.05 | 2.07 | 5.52 | 89.41 | 89.6
5.3 |  | 0.4 | 1 | 1.97 | 4.57 | 89.04 | 89.4
5.4 |  | 0.4 | 1.01 | 1.91 | 4.03 | 88.71 | 89
5.5 |  | 0.4 | 0.95 | 1.78 | 3.79 | 88.34 | 88.7
5.6 |  | 0.4 | 0.89 | 1.77 | 3.22 | 88.03 | 88.3
5.7 |  | 0.4 | 0.9 | 1.71 | 3.14 | 87.69 | 88
5.8 |  | 0.4 | 0.91 | 1.58 | 2.95 | 87.47 | 87.7
5.9 |  | 0.4 | 0.86 | 1.57 | 2.77 | 87.02 | 87.4
6 |  | 0.4 | 0.79 | 1.5 | 2.68 | 86.75 | 87
6.1 |  | 0.29 | 0.8 | 1.47 | 2.64 | 86.32 | 86.7
6.2 |  | 0.3 | 0.81 | 1.4 | 2.25 | 86.05 | 86.4
6.3 |  | 0.3 | 0.76 | 1.41 | 2.3 | 85.62 | 86.1
6.4 |  | 0.3 | 0.69 | 1.28 | 2.27 | 85.34 | 85.8
6.5 |  | 0.3 | 0.7 | 1.3 | 2.17 | 85.05 | 85.5
6.6 |  | 0.3 | 0.7 | 1.27 | 2.07 | 84.63 | 85.1
6.7 |  | 0.3 | 0.7 | 1.2 | 2 | 84.22 | 84.8
6.8 |  | 0.3 | 0.7 | 1.2 | 1.97 | 83.95 | 84.5
6.9 |  | 0.3 | 0.71 | 1.21 | 1.9 | 83.52 | 84.2
7 |  | 0.3 | 0.66 | 1.09 | 1.87 | 83.2 | 83.8
7.1 |  | 0.3 | 0.59 | 1.1 | 1.8 | 82.84 | 83.5
7.2 |  | 0.3 | 0.6 | 1.1 | 1.77 | 82.55 | 83.2
7.3 |  | 0.31 | 0.6 | 1.11 | 1.7 | 82.13 | 82.8
7.4 |  | 0.25 | 0.6 | 0.98 | 1.67 | 81.73 | 82.5
7.5 |  | 0.2 | 0.6 | 1 | 1.6 | 81.32 | 82.2
7.6 |  | 0.2 | 0.6 | 1 | 1.6 | 81.04 | 81.9
7.7 |  | 0.2 | 0.61 | 1 | 1.57 | 80.59 | 81.6
7.8 |  | 0.2 | 0.56 | 0.98 | 1.5 | 80.25 | 81.2
7.9 |  | 0.2 | 0.49 | 0.9 | 1.52 | 79.79 | 80.9
8 |  | 0.2 | 0.5 | 0.9 | 1.35 | 79.42 | 80.64
8.1 |  | 0.2 | 0.5 | 0.9 | 1.4 | 78.9 | 80.3
8.2 |  | 0.2 | 0.5 | 0.9 | 1.4 | 78.41 | 79.9
8.3 |  | 0.2 | 0.5 | 0.88 | 1.38 | 77.76 | 79.6
8.4 |  | 0.2 | 0.5 | 0.8 | 1.3 | 77.01 | 79.3
8.5 |  | 0.2 | 0.5 | 0.8 | 1.3 | 76 | 78.9
8.6 |  | 0.2 | 0.5 | 0.8 | 1.32 | 74.41 | 78.7
8.7 |  | 0.2 | 0.5 | 0.8 | 1.15 | 71.61 | 78.3
8.8 |  |  | 0.5 | 0.8 | 1.2 | 66.87 | 78.04
8.9 |  |  | 0.51 | 0.8 | 1.2 | 60.06 | 77.7
9 |  |  | 0.46 | 0.8 | 1.2 | 50.97 | 77.3
9.1 |  |  | 0.39 | 0.69 | 1.18 | 41.91 | 77.04
9.2 |  |  | 0.4 | 0.7 | 1.1 | 32.03 | 76.7
9.3 |  |  | 0.4 | 0.7 | 1.1 | 23.63 | 76.45
9.4 |  |  | 0.4 | 0.7 | 1.1 | 16.43 | 76.14
9.5 |  |  | 0.4 | 0.7 | 1.08 | 11.35 | 75.84
9.6 |  |  | 0.4 | 0.7 | 1 | 8.56 | 75.54
9.7 |  |  |  | 0.7 | 1 | 7.07 | 75.14
9.8 |  |  |  | 0.71 | 1 | 6.23 | 74.71
9.9 |  |  |  | 0.59 | 1 | 5.59 | 74.42
10 |  |  |  | 0.6 | 1.02 | 5.24 | 74.12
10.1 |  |  |  | 0.6 | 0.89 | 4.94 | 73.85
10.2 |  |  |  | 0.6 | 0.85 | 4.62 | 73.41
10.3 |  |  |  | 0.6 | 0.9 | 4.46 | 73.12
10.4 |  |  |  | 0.6 | 0.9 | 4.3 | 72.82
10.5 |  |  |  | 0.6 | 0.9 | 4.05 | 72.55
10.6 |  |  |  |  | 0.9 | 3.98 | 72.11
10.7 |  |  |  |  | 0.88 | 3.9 | 71.85
10.8 |  |  |  |  | 0.8 | 3.69 | 71.53
10.9 |  |  |  |  | 0.8 | 3.59 | 71.11
11 |  |  |  |  | 0.8 | 3.49 | 70.82
11.1 |  |  |  |  | 0.8 | 3.38 | 70.55
11.2 |  |  |  |  | 0.8 | 3.27 | 70.11
11.3 |  |  |  |  | 0.8 | 3.27 | 69.85
11.4 |  |  |  |  | 0.8 | 3.18 | 69.41
11.5 | [5 empty] | 3.1 | 69.15
11.6 | [5 empty] | 3.08 | 68.74
11.7 | [5 empty] | 3.02 | 68.34
11.8 | [5 empty] | 2.87 | 67.94
11.9 | [5 empty] | 2.87 | 67.45
12 | [5 empty] | 2.8 | 66.95
12.1 | [5 empty] | 2.78 | 66.26
12.2 | [5 empty] | 2.7 | 65.4
12.3 | [5 empty] | 2.68 | 64.06
12.4 | [5 empty] | 2.6 | 62.11
12.5 | [5 empty] | 2.58 | 57.21
12.6 | [5 empty] | 2.5 | 51.15
12.7 | [5 empty] | 2.48 | 44.69
12.8 | [5 empty] | 2.4 | 36.58
12.9 | [5 empty] | 2.4 | 28.59
13 | [5 empty] | 2.38 | 20.89
13.1 | [5 empty] | 2.3 | 14.74
13.2 | [5 empty] | 2.28 | 10.31
13.3 | [5 empty] | 2.2 | 8.15
13.4 | [5 empty] | 2.2 | 6.86
13.5 | [5 empty] | 2.18 | 6.04
13.6 | [5 empty] | 2.1 | 5.57
13.7 | [5 empty] | 2.1 | 5.36
13.8 | [5 empty] | 2.09 | 4.89
13.9 | [5 empty] | 2 | 4.73
14 | [5 empty] | 2 | 4.53
14.1 | [5 empty] | 2.01 | 4.33
14.2 | [5 empty] | 1.88 | 4.22
14.3 | [5 empty] | 1.9 | 4.23
14.4 | [5 empty] | 1.9 | 3.87
14.5 | [5 empty] | 1.89 | 3.81
14.6 | [5 empty] | 1.8 | 3.71
14.7 | [5 empty] | 1.8 | 3.62
14.8 | [5 empty] | 1.8 | 3.72
14.9 | [5 empty] | 1.79 | 3.58
15 | [5 empty] | 1.7 | 3.36
15.1 | [5 empty] | 1.7 | 3.32
15.2 | [5 empty] | 1.7 | 3.42
15.3 | [5 empty] | 1.69 | 3.16
15.4 | [5 empty] | 1.6 | 3.11
15.5 | [5 empty] | 1.6 | 3.1
15.6 | [5 empty] | 1.6 | 3.01
15.7 | [5 empty] | 1.59 | 2.91
15.8 | [5 empty] | 1.5 | 2.9
15.9 | [5 empty] | 1.5 | 2.81
16 | [6 empty] | 2.8
16.1 | [6 empty] | 2.71
16.2 | [6 empty] | 2.7
16.3 | [6 empty] | 2.61
16.4 | [6 empty] | 2.6
16.5 | [6 empty] | 2.71
16.6 | [6 empty] | 2.46
16.7 | [6 empty] | 2.5
16.8 | [6 empty] | 2.41
16.9 | [6 empty] | 2.4
17 | [6 empty] | 2.4
17.1 | [6 empty] | 2.32
17.2 | [6 empty] | 2.3
17.3 | [6 empty] | 2.21
17.4 | [6 empty] | 2.2
17.5 | [6 empty] | 2.2
17.6 | [6 empty] | 2.25
17.7 | [6 empty] | 2.08
17.8 | [6 empty] | 2.1
17.9 | [6 empty] | 2.1
18 | [6 empty] | 2.02
18.1 | [6 empty] | 2
18.2 | [6 empty] | 2
18.3 | [6 empty] | 1.92
18.4 | [6 empty] | 1.9
18.5 | [6 empty] | 1.9
18.6 | [6 empty] | 1.9
18.7 | [6 empty] | 1.95
18.8 | [6 empty] | 1.78
18.9 | [6 empty] | 1.8
19 | [6 empty] | 1.8
19.1 | [6 empty] | 1.8
19.2 | [6 empty] | 1.72
19.3 | [6 empty] | 1.7
19.4 | [6 empty] | 1.7
19.5 | [6 empty] | 1.7
19.6 | [6 empty] | 1.62
19.7
19.8
19.9
20
20.1
20.2
20.3
20.4
20.5
20.6
20.7
20.8
20.9
21
21.1
21.2
21.3
21.4
21.5
21.6
21.7
21.8
21.9
22
22.1
22.2
22.3
22.4
22.5
22.6
22.7
22.8
22.9
23
23.1
23.2
23.3
23.4
23.5
23.6
23.7
23.8
23.9
24
24.1
24.2
24.3
24.4
24.5
24.6
24.7
24.8
24.9
25
25.1
25.2

## Open Field Profiles at 5cm
Lateral Profiles
Open field
SSD = 90 cm
Depth = 5 cm
Scanning direction: Crossline (i.e. along the leaf travel direction)
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 2x2 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 28x28 cm2
-25.2
-25.1
-25
-24.9
-24.8
-24.7
-24.6
-24.5
-24.4
-24.3
-24.2
-24.1
-24
-23.9
-23.8
-23.7
-23.6
-23.5
-23.4
-23.3
-23.2
-23.1
-23
-22.9
-22.8
-22.7
-22.6
-22.5
-22.4
-22.3
-22.2
-22.1
-22
-21.9
-21.8
-21.7
-21.6
-21.5
-21.4
-21.3
-21.2
-21.1
-21
-20.9
-20.8
-20.7
-20.6
-20.5
-20.4
-20.3
-20.2 | [6 empty] | 2.3
-20.1 | [6 empty] | 2.3
-20 | [6 empty] | 2.32
-19.9 | [6 empty] | 2.41
-19.8 | [6 empty] | 2.38
-19.7 | [6 empty] | 2.55
-19.6 | [6 empty] | 2.5
-19.5 | [6 empty] | 2.48
-19.4 | [6 empty] | 2.63
-19.3 | [6 empty] | 2.6
-19.2 | [6 empty] | 2.62
-19.1 | [6 empty] | 2.7
-19 | [6 empty] | 2.68
-18.9 | [6 empty] | 2.83
-18.8 | [6 empty] | 2.8
-18.7 | [6 empty] | 2.82
-18.6 | [6 empty] | 2.9
-18.5 | [6 empty] | 2.92
-18.4 | [6 empty] | 3
-18.3 | [6 empty] | 3.02
-18.2 | [6 empty] | 3.1
-18.1 | [6 empty] | 3.12
-18 | [6 empty] | 3.2
-17.9 | [6 empty] | 3.21
-17.8 | [6 empty] | 3.3
-17.7 | [6 empty] | 3.31
-17.6 | [6 empty] | 3.42
-17.5 | [6 empty] | 3.5
-17.4 | [6 empty] | 3.51
-17.3 | [6 empty] | 3.6
-17.2 | [6 empty] | 3.62
-17.1 | [6 empty] | 3.71
-17 | [6 empty] | 3.79
-16.9 | [6 empty] | 3.91
-16.8 | [6 empty] | 3.91
-16.7 | [6 empty] | 4.01
-16.6 | [6 empty] | 4.11
-16.5 | [6 empty] | 4.21
-16.4 | [5 empty] | 2.1 | 4.29
-16.3 | [5 empty] | 2.1 | 4.42
-16.2 | [5 empty] | 2.13 | 4.41
-16.1 | [5 empty] | 2.21 | 4.6
-16 | [5 empty] | 2.19 | 4.71
-15.9 | [5 empty] | 2.23 | 4.81
-15.8 | [5 empty] | 2.31 | 4.93
-15.7 | [5 empty] | 2.29 | 5.01
-15.6 | [5 empty] | 2.33 | 5.1
-15.5 | [5 empty] | 2.41 | 5.33
-15.4 | [5 empty] | 2.39 | 5.41
-15.3 | [5 empty] | 2.43 | 5.5
-15.2 | [5 empty] | 2.5 | 5.73
-15.1 | [5 empty] | 2.53 | 5.82
-15 | [5 empty] | 2.6 | 6
-14.9 | [5 empty] | 2.63 | 6.22
-14.8 | [5 empty] | 2.71 | 6.42
-14.7 | [5 empty] | 2.61 | 6.63
-14.6 | [5 empty] | 2.92 | 6.92
-14.5 | [5 empty] | 2.76 | 7.13
-14.4 | [5 empty] | 2.81 | 7.44
-14.3 | [5 empty] | 2.94 | 7.73
-14.2 | [5 empty] | 3 | 8.25
-14.1 | [5 empty] | 3.01 | 8.65
-14 | [5 empty] | 3.11 | 9.38
-13.9 | [5 empty] | 3.12 | 10.31
-13.8 | [5 empty] | 3.21 | 11.69
-13.7 | [5 empty] | 3.21 | 14.27
-13.6 | [5 empty] | 3.31 | 18.21
-13.5 | [5 empty] | 3.31 | 24.01
-13.4 | [5 empty] | 3.41 | 30.7
-13.3 | [5 empty] | 3.54 | 37.84
-13.2 | [5 empty] | 3.6 | 44.94
-13.1 | [5 empty] | 3.61 | 51.39
-13 | [5 empty] | 3.71 | 56.1
-12.9 | [5 empty] | 3.81 | 60.13
-12.8 | [5 empty] | 3.9 | 62.19
-12.7 | [5 empty] | 3.91 | 63.41
-12.6 | [5 empty] | 4 | 64.28
-12.5 | [5 empty] | 4.14 | 65.08
-12.4 | [5 empty] | 4.21 | 65.66
-12.3 | [5 empty] | 4.35 | 66.26
-12.2 | [5 empty] | 4.51 | 66.65
-12.1 | [5 empty] | 4.53 | 67.13
-12 | [5 empty] | 4.73 | 67.66
-11.9 | [5 empty] | 4.82 | 68
-11.8 | [5 empty] | 4.92 | 68.44
-11.7 |  |  |  |  | 1.18 | 5.04 | 68.84
-11.6 |  |  |  |  | 1.21 | 5.22 | 69.2
-11.5 |  |  |  |  | 1.2 | 5.34 | 69.63
-11.4 |  |  |  |  | 1.2 | 5.53 | 70.15
-11.3 |  |  |  |  | 1.19 | 5.63 | 70.42
-11.2 |  |  |  |  | 1.27 | 5.84 | 70.85
-11.1 |  |  |  |  | 1.31 | 6.04 | 71.14
-11 |  |  |  |  | 1.3 | 6.24 | 71.53
-10.9 |  |  |  |  | 1.29 | 6.45 | 71.84
-10.8 |  |  |  |  | 1.37 | 6.76 | 72.23
-10.7 |  |  |  | 0.84 | 1.41 | 7.06 | 72.5
-10.6 |  |  |  | 0.9 | 1.4 | 7.36 | 72.84
-10.5 |  |  |  | 0.9 | 1.39 | 7.67 | 73.24
-10.4 |  |  |  | 0.9 | 1.47 | 8.09 | 73.6
-10.3 |  |  |  | 0.9 | 1.51 | 8.61 | 73.94
-10.2 |  |  |  | 0.9 | 1.49 | 9.26 | 74.3
-10.1 |  |  |  | 1 | 1.57 | 10.21 | 74.6
-10 |  |  |  | 1 | 1.61 | 11.52 | 75
-9.9 |  |  |  | 1 | 1.59 | 13.89 | 75.3
-9.8 |  |  | 0.57 | 1 | 1.67 | 17.98 | 75.6
-9.7 |  |  | 0.6 | 1 | 1.71 | 24.08 | 76.03
-9.6 |  |  | 0.6 | 1.1 | 1.69 | 31.68 | 76.4
-9.5 |  |  | 0.6 | 1.1 | 1.78 | 39.89 | 76.6
-9.4 |  |  | 0.47 | 1.1 | 1.8 | 48.35 | 77
-9.3 |  |  | 0.77 | 1.1 | 1.87 | 56.23 | 77.3
-9.2 |  |  | 0.7 | 1.1 | 1.9 | 63.14 | 77.65
-9.1 |  |  | 0.7 | 1.2 | 1.97 | 67.98 | 78
-9 |  |  | 0.7 | 1.2 | 2.01 | 70.95 | 78.3
-8.9 |  |  | 0.7 | 1.2 | 1.99 | 73.03 | 78.6
-8.8 |  | 0.4 | 0.7 | 1.3 | 2.07 | 74.2 | 78.9
-8.7 |  | 0.3 | 0.7 | 1.3 | 2.1 | 75.08 | 79.3
-8.6 |  | 0.4 | 0.77 | 1.3 | 2.16 | 75.87 | 79.6
-8.5 |  | 0.4 | 0.8 | 1.4 | 2.2 | 76.56 | 80.03
-8.4 |  | 0.4 | 0.8 | 1.4 | 2.25 | 77.15 | 80.2
-8.3 |  | 0.4 | 0.8 | 1.4 | 2.37 | 77.65 | 80.6
-8.2 |  | 0.4 | 0.8 | 1.5 | 2.39 | 78.15 | 80.9
-8.1 |  | 0.4 | 0.87 | 1.5 | 2.46 | 78.65 | 81.2
-8 |  | 0.4 | 0.9 | 1.5 | 2.5 | 79.15 | 81.5
-7.9 |  | 0.4 | 0.9 | 1.6 | 2.55 | 79.64 | 81.9
-7.8 |  | 0.4 | 0.9 | 1.6 | 2.66 | 80.05 | 82.2
-7.7 |  | 0.5 | 0.77 | 1.7 | 2.76 | 80.55 | 82.5
-7.6 |  | 0.5 | 1.12 | 1.7 | 2.86 | 80.94 | 82.9
-7.5 |  | 0.5 | 0.98 | 1.8 | 2.96 | 81.34 | 83.2
-7.4 |  | 0.5 | 1 | 1.8 | 3.06 | 81.74 | 83.5
-7.3 |  | 0.5 | 1.06 | 1.9 | 3.16 | 82.14 | 83.8
-7.2 |  | 0.5 | 1.11 | 1.9 | 3.26 | 82.54 | 84.1
-7.1 |  | 0.5 | 1.1 | 2 | 3.36 | 82.94 | 84.4
-7 |  | 0.5 | 1.09 | 2.1 | 3.45 | 83.34 | 84.8
-6.9 |  | 0.6 | 1.16 | 2.1 | 3.62 | 83.73 | 85.2
-6.8 |  | 0.6 | 1.21 | 2.2 | 3.76 | 84.04 | 85.4
-6.7 |  | 0.6 | 1.19 | 2.3 | 3.92 | 84.54 | 85.7
-6.6 |  | 0.6 | 1.26 | 2.3 | 4.06 | 84.83 | 86
-6.5 |  | 0.6 | 1.31 | 2.4 | 4.21 | 85.23 | 86.4
-6.4 |  | 0.6 | 1.29 | 2.5 | 4.42 | 85.54 | 86.7
-6.3 |  | 0.68 | 1.36 | 2.6 | 4.62 | 85.93 | 87
-6.2 |  | 0.7 | 1.4 | 2.7 | 4.81 | 86.23 | 87.4
-6.1 |  | 0.7 | 1.44 | 2.8 | 5.07 | 86.64 | 87.7
-6 |  | 0.7 | 1.57 | 2.9 | 5.38 | 87.03 | 87.9
-5.9 |  | 0.69 | 1.6 | 3.1 | 5.67 | 87.34 | 88.2
-5.8 |  | 0.8 | 1.66 | 3.17 | 6.03 | 87.74 | 88.6
-5.7 |  | 0.8 | 1.7 | 3.3 | 6.49 | 88.03 | 88.9
-5.6 |  | 0.8 | 1.74 | 3.48 | 7.04 | 88.33 | 89.2
-5.5 |  | 0.8 | 1.86 | 3.67 | 7.75 | 88.64 | 89.5
-5.4 |  | 0.88 | 1.9 | 3.79 | 8.71 | 89.03 | 89.8
-5.3 |  | 0.9 | 1.94 | 4 | 10.1 | 89.34 | 90.1
-5.2 |  | 0.9 | 2.05 | 4.28 | 12.57 | 89.73 | 90.4
-5.1 |  | 0.98 | 2.15 | 4.47 | 17.11 | 90.04 | 90.7
-5 |  | 1 | 2.25 | 4.77 | 24.01 | 90.43 | 90.9
-4.9 | 0.3 | 0.99 | 2.34 | 5.15 | 33.08 | 90.62 | 91.2
-4.8 | 0.3 | 1.1 | 2.5 | 5.56 | 43.24 | 90.94 | 91.6
-4.7 | 0.3 | 1.1 | 2.65 | 6.05 | 53.72 | 91.33 | 91.9
-4.6 | 0.3 | 1.18 | 2.8 | 6.72 | 63.94 | 91.53 | 92.1
-4.5 | 0.3 | 1.19 | 2.95 | 7.51 | 72.76 | 91.94 | 92.4
-4.4 | 0.3 | 1.3 | 3.09 | 8.77 | 79.4 | 92.23 | 92.7
-4.3 | 0.3 | 1.29 | 3.3 | 10.78 | 83.71 | 92.53 | 93
-4.2 | 0.3 | 1.4 | 3.5 | 14.43 | 86.27 | 92.83 | 93.2
-4.1 | 0.4 | 1.48 | 3.73 | 20.51 | 87.88 | 93.14 | 93.5
-4 | 0.4 | 1.49 | 4.11 | 29.03 | 89.09 | 93.53 | 93.77
-3.9 | 0.4 | 1.59 | 4.25 | 38.73 | 90 | 93.72 | 94.1
-3.8 | 0.4 | 1.69 | 4.92 | 49.79 | 90.72 | 94.03 | 94.37
-3.7 | 0.4 | 1.79 | 5.37 | 60.79 | 91.37 | 94.3 | 94.6
-3.6 | 0.47 | 1.89 | 6.05 | 70.73 | 91.89 | 94.53 | 94.86
-3.5 | 0.5 | 2.09 | 6.94 | 78.72 | 92.45 | 94.83 | 95.1
-3.4 | 0.5 | 2.15 | 8.47 | 84.57 | 93.02 | 95.1 | 95.4
-3.3 | 0.5 | 2.37 | 11 | 87.8 | 93.36 | 95.3 | 95.56
-3.2 | 0.6 | 2.5 | 15.58 | 90.08 | 93.79 | 95.6 | 95.86
-3.1 | 0.6 | 2.66 | 22.66 | 91.41 | 94.24 | 96 | 96.09
-3 | 0.6 | 2.92 | 31.4 | 92.37 | 94.65 | 96.1 | 96.3
-2.9 | 0.68 | 3.15 | 43.14 | 93.2 | 94.95 | 96.3 | 96.5
-2.8 | 0.69 | 3.53 | 54.82 | 93.83 | 95.31 | 96.5 | 96.66
-2.7 | 0.8 | 3.79 | 65.96 | 94.43 | 95.69 | 96.8 | 96.89
-2.6 | 0.79 | 4.3 | 75.81 | 94.93 | 95.97 | 97 | 97.2
-2.5 | 0.9 | 4.98 | 83.16 | 95.37 | 96.26 | 97.2 | 97.36
-2.4 | 0.96 | 5.85 | 88.36 | 95.74 | 96.54 | 97.4 | 97.6
-2.3 | 1.06 | 7.2 | 91.47 | 96.13 | 96.85 | 97.6 | 97.77
-2.2 | 1.15 | 9.97 | 93.23 | 96.57 | 97.09 | 97.8 | 97.89
-2.1 | 1.3 | 16.79 | 94.35 | 96.83 | 97.34 | 98 | 98.1
-2 | 1.47 | 29.98 | 95.28 | 97.14 | 97.6 | 98.1 | 98.26
-1.9 | 1.61 | 52.97 | 96.03 | 97.46 | 97.79 | 98.3 | 98.5
-1.8 | 1.85 | 71.6 | 96.51 | 97.69 | 97.98 | 98.5 | 98.6
-1.7 | 2.22 | 84.89 | 97.09 | 98.14 | 98.24 | 98.6 | 98.69
-1.6 | 2.68 | 91.22 | 97.25 | 98.18 | 98.51 | 98.9 | 98.9
-1.5 | 3.21 | 94.25 | 97.75 | 98.46 | 98.64 | 98.9 | 99
-1.4 | 4.22 | 95.65 | 98.2 | 98.65 | 98.79 | 99 | 99.1
-1.3 | 6.03 | 96.64 | 98.39 | 98.8 | 98.96 | 99.2 | 99.2
-1.2 | 10.67 | 97.32 | 98.67 | 98.95 | 99.03 | 99.3 | 99.3
-1.1 | 21.36 | 97.75 | 98.87 | 99.26 | 99.19 | 99.4 | 99.47
-1 | 41.27 | 98.15 | 99.11 | 99.29 | 99.35 | 99.5 | 99.5
-0.9 | 64.28 | 98.64 | 99.28 | 99.45 | 99.44 | 99.6 | 99.6
-0.8 | 81.73 | 99.07 | 99.42 | 99.49 | 99.54 | 99.6 | 99.75
-0.7 | 91.45 | 99.39 | 99.56 | 99.66 | 99.64 | 99.7 | 99.8
-0.6 | 95.92 | 99.3 | 99.71 | 99.75 | 99.75 | 99.8 | 99.85
-0.5 | 97.83 | 99.45 | 99.79 | 99.8 | 99.79 | 99.8 | 99.9
-0.4 | 98.69 | 99.74 | 99.86 | 99.84 | 99.83 | 99.9 | 99.9
-0.3 | 99.27 | 99.74 | 99.9 | 100 | 99.95 | 99.9 | 100
-0.2 | 99.75 | 99.85 | 99.96 | 99.95 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 99.95 | 99.99 | 100 | 100
0 | 99.81 | 99.85 | 100 | 100 | 99.99 | 100 | 100
0.1 | 99.98 | 100 | 100 | 99.95 | 99.99 | 100 | 100
0.2 | 99.75 | 99.85 | 99.96 | 99.95 | 100 | 100 | 100
0.3 | 99.28 | 99.74 | 99.89 | 100 | 99.95 | 99.9 | 100
0.4 | 98.69 | 99.74 | 99.85 | 99.84 | 99.83 | 99.9 | 99.9
0.5 | 97.83 | 99.45 | 99.79 | 99.8 | 99.79 | 99.8 | 99.9
0.6 | 95.92 | 99.3 | 99.7 | 99.75 | 99.75 | 99.8 | 99.85
0.7 | 91.45 | 99.39 | 99.55 | 99.66 | 99.64 | 99.7 | 99.8
0.8 | 81.73 | 99.06 | 99.4 | 99.49 | 99.54 | 99.6 | 99.75
0.9 | 64.28 | 98.62 | 99.28 | 99.45 | 99.44 | 99.6 | 99.6
1 | 41.27 | 98.15 | 99.11 | 99.29 | 99.35 | 99.5 | 99.5
1.1 | 21.36 | 97.75 | 98.87 | 99.26 | 99.19 | 99.4 | 99.47
1.2 | 10.67 | 97.26 | 98.65 | 98.95 | 99.03 | 99.3 | 99.3
1.3 | 6.03 | 96.64 | 98.4 | 98.8 | 98.96 | 99.2 | 99.2
1.4 | 4.22 | 95.65 | 98.11 | 98.65 | 98.79 | 99 | 99.1
1.5 | 3.21 | 94.24 | 97.81 | 98.46 | 98.64 | 98.9 | 99
1.6 | 2.68 | 91.19 | 97.25 | 98.19 | 98.51 | 98.9 | 98.9
1.7 | 2.22 | 85.01 | 97.08 | 98.12 | 98.24 | 98.6 | 98.69
1.8 | 1.85 | 70.57 | 96.51 | 97.69 | 97.98 | 98.5 | 98.6
1.9 | 1.61 | 51.46 | 96.03 | 97.46 | 97.79 | 98.3 | 98.5
2 | 1.47 | 29.43 | 95.28 | 97.14 | 97.6 | 98.1 | 98.26
2.1 | 1.3 | 16.77 | 94.33 | 96.83 | 97.34 | 98 | 98.1
2.2 | 1.15 | 10.01 | 93.15 | 96.57 | 97.09 | 97.8 | 97.89
2.3 | 1.06 | 7.25 | 91.46 | 96.09 | 96.85 | 97.6 | 97.77
2.4 | 0.96 | 5.85 | 88.36 | 95.7 | 96.54 | 97.4 | 97.6
2.5 | 0.91 | 4.98 | 83.16 | 95.37 | 96.26 | 97.2 | 97.36
2.6 | 0.79 | 4.3 | 75.81 | 94.93 | 95.97 | 97 | 97.2
2.7 | 0.8 | 3.79 | 65.96 | 94.43 | 95.69 | 96.8 | 96.89
2.8 | 0.69 | 3.53 | 54.82 | 93.83 | 95.31 | 96.5 | 96.66
2.9 | 0.68 | 3.15 | 43.13 | 93.2 | 94.95 | 96.3 | 96.5
3 | 0.6 | 2.92 | 31.17 | 92.37 | 94.65 | 96.1 | 96.3
3.1 | 0.6 | 2.66 | 22.56 | 91.41 | 94.24 | 96 | 96.09
3.2 | 0.6 | 2.5 | 15.58 | 90.08 | 93.79 | 95.6 | 95.86
3.3 | 0.5 | 2.35 | 11 | 87.8 | 93.36 | 95.3 | 95.56
3.4 | 0.5 | 2.15 | 8.47 | 84.57 | 93.02 | 95.1 | 95.37
3.5 | 0.5 | 2.08 | 6.88 | 78.72 | 92.45 | 94.8 | 95.1
3.6 | 0.47 | 1.9 | 6.03 | 70.73 | 91.89 | 94.5 | 94.86
3.7 | 0.4 | 1.79 | 5.39 | 60.79 | 91.37 | 94.3 | 94.56
3.8 | 0.4 | 1.69 | 4.84 | 49.79 | 90.72 | 94.03 | 94.37
3.9 | 0.4 | 1.59 | 4.39 | 38.73 | 90 | 93.7 | 94.07
4 | 0.4 | 1.49 | 4.05 | 29.03 | 89.09 | 93.5 | 93.76
4.1 | 0.4 | 1.48 | 3.74 | 20.51 | 87.88 | 93.14 | 93.5
4.2 | 0.3 | 1.4 | 3.49 | 14.43 | 86.27 | 92.83 | 93.2
4.3 | 0.3 | 1.29 | 3.31 | 10.78 | 83.71 | 92.5 | 93
4.4 | 0.3 | 1.3 | 3.09 | 8.77 | 79.4 | 92.23 | 92.67
4.5 | 0.3 | 1.19 | 2.95 | 7.51 | 72.76 | 91.9 | 92.4
4.6 | 0.3 | 1.18 | 2.8 | 6.72 | 63.94 | 91.53 | 92.1
4.7 | 0.3 | 1.1 | 2.65 | 6.05 | 53.72 | 91.3 | 91.9
4.8 | 0.3 | 1.1 | 2.5 | 5.56 | 43.24 | 90.94 | 91.6
4.9 |  | 0.99 | 2.34 | 5.15 | 33.08 | 90.62 | 91.2
5 |  | 1 | 2.25 | 4.77 | 24.01 | 90.43 | 90.9
5.1 |  | 0.98 | 2.15 | 4.47 | 17.11 | 90.04 | 90.7
5.2 |  | 0.9 | 2.05 | 4.28 | 12.57 | 89.73 | 90.4
5.3 |  | 0.9 | 1.94 | 4 | 10.1 | 89.34 | 90.1
5.4 |  | 0.88 | 1.9 | 3.79 | 8.71 | 89.03 | 89.8
5.5 |  | 0.8 | 1.86 | 3.67 | 7.75 | 88.64 | 89.5
5.6 |  | 0.8 | 1.74 | 3.48 | 7.04 | 88.33 | 89.2
5.7 |  | 0.8 | 1.7 | 3.3 | 6.49 | 88.03 | 88.9
5.8 |  | 0.8 | 1.66 | 3.17 | 6.03 | 87.74 | 88.6
5.9 |  | 0.69 | 1.6 | 3.08 | 5.67 | 87.34 | 88.2
6 |  | 0.7 | 1.57 | 2.9 | 5.38 | 87.03 | 87.9
6.1 |  | 0.7 | 1.44 | 2.8 | 5.07 | 86.64 | 87.7
6.2 |  | 0.7 | 1.4 | 2.7 | 4.81 | 86.23 | 87.4
6.3 |  | 0.68 | 1.36 | 2.6 | 4.62 | 85.93 | 87
6.4 |  | 0.6 | 1.29 | 2.5 | 4.42 | 85.54 | 86.7
6.5 |  | 0.6 | 1.31 | 2.4 | 4.21 | 85.23 | 86.4
6.6 |  | 0.6 | 1.26 | 2.29 | 4.06 | 84.83 | 86
6.7 |  | 0.6 | 1.19 | 2.28 | 3.92 | 84.54 | 85.7
6.8 |  | 0.6 | 1.21 | 2.2 | 3.76 | 84.04 | 85.4
6.9 |  | 0.6 | 1.16 | 2.1 | 3.62 | 83.73 | 85.2
7 |  | 0.5 | 1.09 | 2.08 | 3.45 | 83.34 | 84.8
7.1 |  | 0.5 | 1.1 | 2 | 3.36 | 82.94 | 84.4
7.2 |  | 0.5 | 1.11 | 1.9 | 3.26 | 82.54 | 84.1
7.3 |  | 0.5 | 1.06 | 1.9 | 3.16 | 82.14 | 83.8
7.4 |  | 0.5 | 1 | 1.8 | 3.06 | 81.74 | 83.5
7.5 |  | 0.5 | 0.98 | 1.8 | 2.96 | 81.34 | 83.2
7.6 |  | 0.5 | 1.12 | 1.7 | 2.86 | 80.94 | 82.9
7.7 |  | 0.5 | 0.77 | 1.7 | 2.76 | 80.55 | 82.5
7.8 |  | 0.4 | 0.9 | 1.6 | 2.66 | 80.05 | 82.2
7.9 |  | 0.4 | 0.9 | 1.6 | 2.55 | 79.64 | 81.9
8 |  | 0.4 | 0.9 | 1.5 | 2.5 | 79.15 | 81.5
8.1 |  | 0.4 | 0.87 | 1.5 | 2.46 | 78.65 | 81.2
8.2 |  | 0.4 | 0.8 | 1.5 | 2.39 | 78.15 | 80.9
8.3 |  | 0.4 | 0.8 | 1.4 | 2.37 | 77.65 | 80.6
8.4 |  | 0.4 | 0.8 | 1.4 | 2.25 | 77.15 | 80.2
8.5 |  | 0.4 | 0.8 | 1.4 | 2.2 | 76.56 | 80.03
8.6 |  | 0.4 | 0.77 | 1.3 | 2.16 | 75.88 | 79.6
8.7 |  | 0.3 | 0.7 | 1.3 | 2.1 | 75.08 | 79.3
8.8 |  |  | 0.7 | 1.3 | 2.07 | 74.21 | 78.9
8.9 |  |  | 0.7 | 1.2 | 1.99 | 72.86 | 78.6
9 |  |  | 0.7 | 1.2 | 2.01 | 70.94 | 78.3
9.1 |  |  | 0.7 | 1.2 | 1.96 | 67.99 | 78
9.2 |  |  | 0.7 | 1.1 | 1.9 | 63.14 | 77.65
9.3 |  |  | 0.77 | 1.1 | 1.86 | 56.23 | 77.3
9.4 |  |  | 0.47 | 1.1 | 1.8 | 48.35 | 77
9.5 |  |  | 0.6 | 1.1 | 1.78 | 39.89 | 76.6
9.6 |  |  | 0.6 | 1.1 | 1.69 | 31.68 | 76.4
9.7 |  |  | 0.6 | 1 | 1.71 | 24.08 | 76.03
9.8 |  |  |  | 1 | 1.67 | 17.98 | 75.6
9.9 |  |  |  | 1 | 1.59 | 13.89 | 75.3
10 |  |  |  | 1 | 1.61 | 11.52 | 75
10.1 |  |  |  | 1 | 1.57 | 10.21 | 74.6
10.2 |  |  |  | 0.9 | 1.49 | 9.26 | 74.3
10.3 |  |  |  | 0.9 | 1.51 | 8.61 | 73.94
10.4 |  |  |  | 0.9 | 1.47 | 8.09 | 73.6
10.5 |  |  |  | 0.9 | 1.39 | 7.67 | 73.24
10.6 |  |  |  | 0.9 | 1.4 | 7.36 | 72.84
10.7 |  |  |  |  | 1.41 | 7.06 | 72.5
10.8 |  |  |  |  | 1.37 | 6.76 | 72.23
10.9 |  |  |  |  | 1.29 | 6.45 | 71.84
11 |  |  |  |  | 1.3 | 6.24 | 71.53
11.1 |  |  |  |  | 1.31 | 6.04 | 71.14
11.2 |  |  |  |  | 1.27 | 5.84 | 70.85
11.3 |  |  |  |  | 1.19 | 5.63 | 70.42
11.4 |  |  |  |  | 1.2 | 5.53 | 70.15
11.5 |  |  |  |  | 1.2 | 5.34 | 69.63
11.6 |  |  |  |  | 1.21 | 5.22 | 69.2
11.7 | [5 empty] | 5.04 | 68.84
11.8 | [5 empty] | 4.92 | 68.44
11.9 | [5 empty] | 4.82 | 68
12 | [5 empty] | 4.73 | 67.66
12.1 | [5 empty] | 4.52 | 67.13
12.2 | [5 empty] | 4.51 | 66.65
12.3 | [5 empty] | 4.32 | 66.26
12.4 | [5 empty] | 4.21 | 65.66
12.5 | [5 empty] | 4.14 | 65.08
12.6 | [5 empty] | 4 | 64.28
12.7 | [5 empty] | 3.91 | 63.41
12.8 | [5 empty] | 3.9 | 62.19
12.9 | [5 empty] | 3.81 | 60.13
13 | [5 empty] | 3.71 | 56.1
13.1 | [5 empty] | 3.61 | 51.39
13.2 | [5 empty] | 3.6 | 44.94
13.3 | [5 empty] | 3.54 | 37.84
13.4 | [5 empty] | 3.41 | 30.7
13.5 | [5 empty] | 3.31 | 24.01
13.6 | [5 empty] | 3.31 | 18.21
13.7 | [5 empty] | 3.21 | 14.27
13.8 | [5 empty] | 3.21 | 11.69
13.9 | [5 empty] | 3.12 | 10.31
14 | [5 empty] | 3.11 | 9.38
14.1 | [5 empty] | 3.01 | 8.65
14.2 | [5 empty] | 3 | 8.25
14.3 | [5 empty] | 2.94 | 7.73
14.4 | [5 empty] | 2.81 | 7.44
14.5 | [5 empty] | 2.76 | 7.13
14.6 | [5 empty] | 2.92 | 6.92
14.7 | [5 empty] | 2.61 | 6.63
14.8 | [5 empty] | 2.71 | 6.42
14.9 | [5 empty] | 2.63 | 6.22
15 | [5 empty] | 2.6 | 6
15.1 | [5 empty] | 2.53 | 5.82
15.2 | [5 empty] | 2.5 | 5.73
15.3 | [5 empty] | 2.43 | 5.5
15.4 | [5 empty] | 2.39 | 5.41
15.5 | [5 empty] | 2.41 | 5.33
15.6 | [5 empty] | 2.33 | 5.1
15.7 | [5 empty] | 2.29 | 5.01
15.8 | [5 empty] | 2.31 | 4.93
15.9 | [5 empty] | 2.23 | 4.81
16 | [5 empty] | 2.19 | 4.71
16.1 | [5 empty] | 2.21 | 4.6
16.2 | [5 empty] | 2.13 | 4.41
16.3 | [5 empty] | 2.1 | 4.42
16.4 | [6 empty] | 4.29
16.5 | [6 empty] | 4.21
16.6 | [6 empty] | 4.11
16.7 | [6 empty] | 4.01
16.8 | [6 empty] | 3.91
16.9 | [6 empty] | 3.91
17 | [6 empty] | 3.79
17.1 | [6 empty] | 3.71
17.2 | [6 empty] | 3.62
17.3 | [6 empty] | 3.6
17.4 | [6 empty] | 3.51
17.5 | [6 empty] | 3.5
17.6 | [6 empty] | 3.42
17.7 | [6 empty] | 3.31
17.8 | [6 empty] | 3.3
17.9 | [6 empty] | 3.21
18 | [6 empty] | 3.2
18.1 | [6 empty] | 3.12
18.2 | [6 empty] | 3.1
18.3 | [6 empty] | 3.02
18.4 | [6 empty] | 3
18.5 | [6 empty] | 2.92
18.6 | [6 empty] | 2.9
18.7 | [6 empty] | 2.82
18.8 | [6 empty] | 2.8
18.9 | [6 empty] | 2.83
19 | [6 empty] | 2.68
19.1 | [6 empty] | 2.7
19.2 | [6 empty] | 2.62
19.3 | [6 empty] | 2.6
19.4 | [6 empty] | 2.63
19.5 | [6 empty] | 2.48
19.6 | [6 empty] | 2.5
19.7 | [6 empty] | 2.55
19.8 | [6 empty] | 2.38
19.9 | [6 empty] | 2.41
20 | [6 empty] | 2.32
20.1 | [6 empty] | 2.3
20.2
20.3
20.4
20.5
20.6
20.7
20.8
20.9
21
21.1
21.2
21.3
21.4
21.5
21.6
21.7
21.8
21.9
22
22.1
22.2
22.3
22.4
22.5
22.6
22.7
22.8
22.9
23
23.1
23.2
23.3
23.4
23.5
23.6
23.7
23.8
23.9
24
24.1
24.2
24.3
24.4
24.5
24.6
24.7
24.8
24.9
25
25.1
25.2

## Open Field Profiles at 10cm
Lateral Profiles
Open field
SSD = 90 cm
Depth=10cm
Scanning direction: Crossline (i.e. along the leaf travel direction)
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 2x2 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 28x28 cm2
-25.2
-25.1
-25
-24.9
-24.8
-24.7
-24.6
-24.5
-24.4
-24.3
-24.2
-24.1
-24
-23.9
-23.8
-23.7
-23.6
-23.5
-23.4
-23.3
-23.2
-23.1
-23
-22.9
-22.8
-22.7
-22.6
-22.5
-22.4
-22.3
-22.2
-22.1
-22
-21.9
-21.8
-21.7
-21.6
-21.5
-21.4
-21.3
-21.2
-21.1
-21
-20.9
-20.8 | [6 empty] | 3.5
-20.7 | [6 empty] | 3.52
-20.6 | [6 empty] | 3.59
-20.5 | [6 empty] | 3.71
-20.4 | [6 empty] | 3.7
-20.3 | [6 empty] | 3.81
-20.2 | [6 empty] | 3.82
-20.1 | [6 empty] | 3.89
-20 | [6 empty] | 4.01
-19.9 | [6 empty] | 4.02
-19.8 | [6 empty] | 4.09
-19.7 | [6 empty] | 4.21
-19.6 | [6 empty] | 4.2
-19.5 | [6 empty] | 4.3
-19.4 | [6 empty] | 4.41
-19.3 | [6 empty] | 4.42
-19.2 | [6 empty] | 4.51
-19.1 | [6 empty] | 4.59
-19 | [6 empty] | 4.71
-18.9 | [6 empty] | 4.72
-18.8 | [6 empty] | 4.82
-18.7 | [6 empty] | 4.92
-18.6 | [6 empty] | 5.01
-18.5 | [6 empty] | 5.09
-18.4 | [6 empty] | 5.2
-18.3 | [6 empty] | 5.3
-18.2 | [6 empty] | 5.41
-18.1 | [6 empty] | 5.5
-18 | [6 empty] | 5.61
-17.9 | [6 empty] | 5.7
-17.8 | [6 empty] | 5.8
-17.7 | [6 empty] | 5.9
-17.6 | [6 empty] | 6
-17.5 | [6 empty] | 6.1
-17.4 | [6 empty] | 6.3
-17.3 | [6 empty] | 6.4
-17.2 | [6 empty] | 6.5
-17.1 | [6 empty] | 6.6
-17 | [6 empty] | 6.7
-16.9 | [5 empty] | 3.07 | 6.9
-16.8 | [5 empty] | 3.25 | 7
-16.7 | [5 empty] | 3.3 | 7.2
-16.6 | [5 empty] | 3.33 | 7.4
-16.5 | [5 empty] | 3.4 | 7.5
-16.4 | [5 empty] | 3.43 | 7.7
-16.3 | [5 empty] | 3.51 | 7.9
-16.2 | [5 empty] | 3.52 | 8
-16.1 | [5 empty] | 3.64 | 8.2
-16 | [5 empty] | 3.7 | 8.4
-15.9 | [5 empty] | 3.73 | 8.6
-15.8 | [5 empty] | 3.81 | 8.9
-15.7 | [5 empty] | 3.82 | 9.1
-15.6 | [5 empty] | 3.93 | 9.3
-15.5 | [5 empty] | 4.03 | 9.6
-15.4 | [5 empty] | 4.1 | 9.8
-15.3 | [5 empty] | 4.12 | 10.1
-15.2 | [5 empty] | 4.24 | 10.5
-15.1 | [5 empty] | 4.3 | 10.8
-15 | [5 empty] | 4.31 | 11.2
-14.9 | [5 empty] | 4.43 | 11.6
-14.8 | [5 empty] | 4.53 | 12.2
-14.7 | [5 empty] | 4.63 | 12.9
-14.6 | [5 empty] | 4.7 | 13.8
-14.5 | [5 empty] | 4.72 | 15.2
-14.4 | [5 empty] | 4.82 | 17.6
-14.3 | [5 empty] | 4.96 | 21.3
-14.2 | [5 empty] | 5.11 | 26.2
-14.1 | [5 empty] | 5.11 | 32.1
-14 | [5 empty] | 5.23 | 38.3
-13.9 | [5 empty] | 5.35 | 44.4
-13.8 | [5 empty] | 5.53 | 49.9
-13.7 | [5 empty] | 5.62 | 54.4
-13.6 | [5 empty] | 5.72 | 57.7
-13.5 | [5 empty] | 5.83 | 59.8
-13.4 | [5 empty] | 5.92 | 61.1
-13.3 | [5 empty] | 6.05 | 62
-13.2 | [5 empty] | 6.22 | 62.73
-13.1 | [5 empty] | 6.34 | 63.4
-13 | [5 empty] | 6.52 | 63.95
-12.9 | [5 empty] | 6.63 | 64.44
-12.8 | [5 empty] | 6.83 | 65.05
-12.7 | [5 empty] | 6.93 | 65.5
-12.6 | [5 empty] | 7.14 | 65.95
-12.5 | [5 empty] | 7.34 | 66.46
-12.4 | [5 empty] | 7.52 | 66.84
-12.3 | [5 empty] | 7.63 | 67.25
-12.2 | [5 empty] | 7.84 | 67.77
-12.1 | [5 empty] | 8.04 | 68.18
-12 | [5 empty] | 8.24 | 68.54
-11.9 |  |  |  |  | 1.94 | 8.46 | 68.96
-11.8 |  |  |  |  | 2 | 8.74 | 69.38
-11.7 |  |  |  |  | 2 | 8.96 | 69.74
-11.6 |  |  |  |  | 2.1 | 9.25 | 70.07
-11.5 |  |  |  |  | 2.1 | 9.45 | 70.46
-11.4 |  |  |  |  | 2.1 | 9.76 | 70.86
-11.3 |  |  |  |  | 2.2 | 10.06 | 71.27
-11.2 |  |  |  |  | 2.2 | 10.37 | 71.66
-11.1 |  |  |  |  | 2.2 | 10.78 | 72.08
-11 |  |  |  |  | 2.3 | 11.19 | 72.35
-10.9 |  |  |  | 1.43 | 2.3 | 11.7 | 72.76
-10.8 |  |  |  | 1.5 | 2.4 | 12.23 | 73.08
-10.7 |  |  |  | 1.5 | 2.4 | 12.94 | 73.47
-10.6 |  |  |  | 1.5 | 2.5 | 13.82 | 73.75
-10.5 |  |  |  | 1.6 | 2.5 | 15.21 | 74.08
-10.4 |  |  |  | 1.6 | 2.6 | 17.34 | 74.46
-10.3 |  |  |  | 1.6 | 2.6 | 21.08 | 74.86
-10.2 |  |  |  | 1.7 | 2.7 | 26.38 | 75.28
-10.1 |  |  |  | 1.7 | 2.8 | 33.03 | 75.56
-10 |  |  |  | 1.7 | 2.8 | 40.51 | 75.96
-9.9 |  |  | 1.03 | 1.8 | 2.9 | 47.97 | 76.28
-9.8 |  |  | 1.1 | 1.8 | 2.9 | 55.08 | 76.67
-9.7 |  |  | 1.1 | 1.9 | 3 | 61.18 | 76.96
-9.6 |  |  | 1.1 | 1.9 | 3.1 | 65.56 | 77.26
-9.5 |  |  | 1.1 | 1.9 | 3.2 | 68.62 | 77.67
-9.4 |  |  | 1.2 | 2 | 3.3 | 70.62 | 77.96
-9.3 |  |  | 1.2 | 2 | 3.3 | 71.8 | 78.37
-9.2 |  |  | 1.2 | 2.1 | 3.4 | 72.69 | 78.66
-9.1 |  |  | 1.19 | 2.2 | 3.5 | 73.56 | 79.07
-9 |  |  | 1.3 | 2.2 | 3.6 | 74.27 | 79.36
-8.9 |  | 0.67 | 1.3 | 2.3 | 3.7 | 74.93 | 79.66
-8.8 |  | 0.7 | 1.31 | 2.3 | 3.7 | 75.56 | 80.06
-8.7 |  | 0.7 | 1.4 | 2.4 | 3.8 | 76.05 | 80.37
-8.6 |  | 0.7 | 1.4 | 2.4 | 4 | 76.56 | 80.68
-8.5 |  | 0.7 | 1.4 | 2.5 | 4 | 77.16 | 80.98
-8.4 |  | 0.7 | 1.5 | 2.6 | 4.2 | 77.64 | 81.28
-8.3 |  | 0.77 | 1.5 | 2.6 | 4.3 | 78.04 | 81.58
-8.2 |  | 0.78 | 1.5 | 2.7 | 4.4 | 78.51 | 81.96
-8.1 |  | 0.8 | 1.6 | 2.8 | 4.5 | 79.05 | 82.27
-8 |  | 0.8 | 1.6 | 2.8 | 4.6 | 79.44 | 82.58
-7.9 |  | 0.8 | 1.7 | 2.9 | 4.8 | 79.9 | 82.97
-7.8 |  | 0.85 | 1.7 | 3 | 4.9 | 80.34 | 83.26
-7.7 |  | 0.9 | 1.7 | 3.1 | 5.1 | 80.75 | 83.56
-7.6 |  | 0.9 | 1.8 | 3.2 | 5.19 | 81.29 | 83.97
-7.5 |  | 0.9 | 1.8 | 3.3 | 5.4 | 81.63 | 84.17
-7.4 |  | 0.95 | 1.9 | 3.4 | 5.57 | 82.05 | 84.47
-7.3 |  | 1 | 2 | 3.5 | 5.69 | 82.54 | 84.88
-7.2 |  | 1 | 2 | 3.6 | 5.88 | 82.84 | 85.18
-7.1 |  | 1 | 2 | 3.7 | 6.08 | 83.34 | 85.49
-7 |  | 1.06 | 2.1 | 3.8 | 6.28 | 83.63 | 85.76
-6.9 |  | 1.1 | 2.2 | 3.9 | 6.47 | 84.04 | 86.17
-6.8 |  | 1.1 | 2.2 | 4.1 | 6.78 | 84.44 | 86.37
-6.7 |  | 1.1 | 2.3 | 4.2 | 6.96 | 84.83 | 86.68
-6.6 |  | 1.15 | 2.4 | 4.4 | 7.28 | 85.14 | 86.98
-6.5 |  | 1.2 | 2.5 | 4.5 | 7.56 | 85.64 | 87.37
-6.4 |  | 1.25 | 2.5 | 4.7 | 7.77 | 85.94 | 87.67
-6.3 |  | 1.25 | 2.6 | 4.8 | 8.15 | 86.44 | 87.95
-6.2 |  | 1.3 | 2.7 | 5 | 8.55 | 86.72 | 88.26
-6.1 |  | 1.3 | 2.8 | 5.2 | 8.94 | 87.04 | 88.47
-6 |  | 1.34 | 2.9 | 5.4 | 9.45 | 87.43 | 88.87
-5.9 |  | 1.4 | 3 | 5.6 | 10.03 | 87.73 | 89.07
-5.8 |  | 1.44 | 3.1 | 5.8 | 10.69 | 88.14 | 89.47
-5.7 |  | 1.5 | 3.2 | 6 | 11.59 | 88.53 | 89.76
-5.6 |  | 1.54 | 3.3 | 6.3 | 12.84 | 88.83 | 89.97
-5.5 |  | 1.6 | 3.4 | 6.57 | 14.8 | 89.14 | 90.27
-5.4 |  | 1.63 | 3.6 | 6.87 | 18.06 | 89.54 | 90.58
-5.3 |  | 1.69 | 3.7 | 7.17 | 23.43 | 89.83 | 90.85
-5.2 |  | 1.81 | 3.9 | 7.47 | 30.62 | 90.13 | 91.17
-5.1 |  | 1.83 | 4 | 7.95 | 39.39 | 90.43 | 91.47
-5 |  | 1.89 | 4.2 | 8.45 | 48.91 | 90.73 | 91.69
-4.9 | 0.5 | 2.01 | 4.37 | 8.94 | 58.46 | 91.04 | 91.98
-4.8 | 0.5 | 2.01 | 4.6 | 9.63 | 67.24 | 91.44 | 92.36
-4.7 | 0.6 | 2.09 | 4.8 | 10.61 | 74.38 | 91.73 | 92.57
-4.6 | 0.6 | 2.2 | 5 | 11.95 | 79.63 | 92.03 | 92.77
-4.5 | 0.6 | 2.3 | 5.2 | 14.08 | 83.07 | 92.33 | 93.07
-4.4 | 0.6 | 2.4 | 5.5 | 17.55 | 85.26 | 92.63 | 93.38
-4.3 | 0.6 | 2.5 | 5.8 | 23.24 | 86.62 | 92.93 | 93.57
-4.2 | 0.6 | 2.6 | 6.1 | 31.13 | 87.83 | 93.23 | 93.88
-4.1 | 0.7 | 2.72 | 6.5 | 40.37 | 88.67 | 93.53 | 94.07
-4 | 0.7 | 2.82 | 7 | 50.4 | 89.47 | 93.83 | 94.35
-3.9 | 0.7 | 3.05 | 7.5 | 60.33 | 90.09 | 94.12 | 94.67
-3.8 | 0.8 | 3.13 | 8.2 | 70.02 | 90.7 | 94.32 | 94.86
-3.7 | 0.8 | 3.25 | 9.01 | 77.33 | 91.27 | 94.6 | 95.08
-3.6 | 0.8 | 3.51 | 10.4 | 82.75 | 91.79 | 94.83 | 95.32
-3.5 | 0.9 | 3.53 | 12.6 | 86.21 | 92.29 | 95.12 | 95.56
-3.4 | 0.9 | 3.7 | 16.2 | 88.4 | 92.71 | 95.3 | 95.76
-3.3 | 1 | 4.02 | 21.44 | 89.86 | 93.19 | 95.6 | 95.96
-3.2 | 1 | 4.33 | 30 | 91 | 93.62 | 95.8 | 96.15
-3.1 | 1.1 | 4.53 | 38.72 | 91.9 | 94.1 | 96.1 | 96.47
-3 | 1.05 | 4.83 | 50.4 | 92.63 | 94.39 | 96.3 | 96.66
-2.9 | 1.2 | 5.24 | 60 | 93.23 | 94.79 | 96.5 | 96.86
-2.8 | 1.16 | 5.65 | 70.1 | 93.76 | 95.19 | 96.8 | 96.99
-2.7 | 1.3 | 6.14 | 78.63 | 94.34 | 95.51 | 97 | 97.27
-2.6 | 1.38 | 6.87 | 85.2 | 94.72 | 95.82 | 97.2 | 97.46
-2.5 | 1.5 | 7.8 | 88.9 | 95.16 | 96.18 | 97.4 | 97.6
-2.4 | 1.58 | 9.31 | 91.24 | 95.63 | 96.51 | 97.6 | 97.76
-2.3 | 1.77 | 12.54 | 92.8 | 95.96 | 96.7 | 97.8 | 97.96
-2.2 | 1.91 | 19.58 | 93.95 | 96.36 | 97.01 | 97.9 | 98.16
-2.1 | 2 | 34.06 | 94.75 | 96.66 | 97.3 | 98.1 | 98.29
-2 | 2.27 | 53.37 | 95.45 | 96.96 | 97.54 | 98.3 | 98.58
-1.9 | 2.45 | 71.89 | 96.05 | 97.36 | 97.71 | 98.5 | 98.59
-1.8 | 2.89 | 83.43 | 96.55 | 97.6 | 98.01 | 98.5 | 98.77
-1.7 | 3.22 | 90.22 | 97.05 | 97.76 | 98.21 | 98.7 | 98.85
-1.6 | 3.85 | 93.31 | 97.29 | 98.06 | 98.41 | 98.9 | 99
-1.5 | 4.68 | 95.2 | 97.7 | 98.36 | 98.61 | 99 | 99.1
-1.4 | 6.33 | 96.3 | 98 | 98.56 | 98.8 | 99.1 | 99.2
-1.3 | 9.67 | 97.1 | 98.3 | 98.76 | 98.94 | 99.2 | 99.36
-1.2 | 17.78 | 97.2 | 98.6 | 98.89 | 99.04 | 99.3 | 99.5
-1.1 | 33.27 | 98 | 98.9 | 99.06 | 99.21 | 99.4 | 99.5
-1 | 53.96 | 98.3 | 99 | 99.19 | 99.33 | 99.5 | 99.55
-0.9 | 73.48 | 98.7 | 99.2 | 99.34 | 99.44 | 99.6 | 99.64
-0.8 | 86.34 | 99 | 99.4 | 99.49 | 99.54 | 99.7 | 99.8
-0.7 | 93.39 | 99 | 99.6 | 99.66 | 99.69 | 99.8 | 99.8
-0.6 | 96.53 | 99.6 | 99.7 | 99.79 | 99.73 | 99.8 | 99.85
-0.5 | 98.35 | 99.6 | 99.7 | 99.9 | 99.84 | 99.9 | 99.9
-0.4 | 99.11 | 99.7 | 99.8 | 99.89 | 99.84 | 100 | 99.95
-0.3 | 99.74 | 100 | 99.9 | 99.89 | 99.84 | 99.9 | 100
-0.2 | 100 | 100 | 99.9 | 99.89 | 100 | 99.9 | 100
-0.1 | 100 | 99.5 | 100 | 100 | 99.93 | 100 | 100
0 | 100 | 99.9 | 100 | 99.99 | 99.94 | 100 | 100
0.1 | 100 | 99.5 | 100 | 100 | 99.93 | 100 | 100
0.2 | 100 | 100 | 99.9 | 99.89 | 100 | 99.9 | 100
0.3 | 99.74 | 100 | 99.9 | 99.89 | 99.84 | 99.9 | 100
0.4 | 99.11 | 99.7 | 99.8 | 99.89 | 99.84 | 100 | 99.95
0.5 | 98.35 | 99.6 | 99.7 | 99.9 | 99.84 | 99.9 | 99.9
0.6 | 96.65 | 99.6 | 99.7 | 99.79 | 99.73 | 99.8 | 99.85
0.7 | 93.15 | 99 | 99.6 | 99.64 | 99.69 | 99.8 | 99.8
0.8 | 85.42 | 99 | 99.4 | 99.49 | 99.54 | 99.7 | 99.8
0.9 | 71.83 | 98.7 | 99.2 | 99.34 | 99.44 | 99.6 | 99.64
1 | 53.96 | 98.3 | 99 | 99.19 | 99.33 | 99.5 | 99.55
1.1 | 33.27 | 98 | 98.9 | 99.04 | 99.21 | 99.4 | 99.5
1.2 | 17.78 | 97.2 | 98.6 | 98.89 | 99.04 | 99.3 | 99.5
1.3 | 9.67 | 97.1 | 98.3 | 98.74 | 98.94 | 99.2 | 99.36
1.4 | 6.33 | 96.3 | 98 | 98.56 | 98.8 | 99.1 | 99.2
1.5 | 4.68 | 95.2 | 97.7 | 98.35 | 98.61 | 99 | 99.1
1.6 | 3.85 | 93.31 | 97.29 | 98.06 | 98.41 | 98.9 | 99
1.7 | 3.22 | 90.22 | 97.05 | 97.76 | 98.21 | 98.7 | 98.85
1.8 | 2.89 | 83.43 | 96.55 | 97.6 | 98.01 | 98.5 | 98.77
1.9 | 2.45 | 71.89 | 96.05 | 97.36 | 97.71 | 98.5 | 98.59
2 | 2.27 | 53.37 | 95.45 | 96.96 | 97.54 | 98.3 | 98.56
2.1 | 2 | 34.06 | 94.75 | 96.66 | 97.3 | 98.1 | 98.29
2.2 | 1.91 | 19.58 | 93.95 | 96.36 | 97.01 | 97.9 | 98.16
2.3 | 1.77 | 12.54 | 92.8 | 95.96 | 96.7 | 97.8 | 97.96
2.4 | 1.58 | 9.31 | 91.26 | 95.63 | 96.51 | 97.6 | 97.76
2.5 | 1.5 | 7.8 | 88.63 | 95.16 | 96.18 | 97.4 | 97.6
2.6 | 1.38 | 6.87 | 85.2 | 94.72 | 95.81 | 97.2 | 97.46
2.7 | 1.3 | 6.14 | 78.63 | 94.34 | 95.51 | 97 | 97.27
2.8 | 1.18 | 5.65 | 70.11 | 93.76 | 95.19 | 96.8 | 96.99
2.9 | 1.2 | 5.24 | 60.05 | 93.23 | 94.79 | 96.5 | 96.86
3 | 1.06 | 4.83 | 49.35 | 92.63 | 94.39 | 96.3 | 96.66
3.1 | 1.1 | 4.53 | 38.7 | 91.9 | 94.1 | 96.1 | 96.47
3.2 | 1 | 4.33 | 29.39 | 91 | 93.62 | 95.8 | 96.15
3.3 | 1 | 4.02 | 21.21 | 89.86 | 93.19 | 95.6 | 95.96
3.4 | 0.9 | 3.7 | 16.2 | 88.4 | 92.71 | 95.3 | 95.76
3.5 | 0.9 | 3.53 | 12.29 | 86.21 | 92.29 | 95.12 | 95.56
3.6 | 0.8 | 3.51 | 10.4 | 82.75 | 91.79 | 94.83 | 95.32
3.7 | 0.8 | 3.25 | 9.02 | 77.31 | 91.27 | 94.6 | 95.08
3.8 | 0.8 | 3.13 | 8.12 | 69.75 | 90.7 | 94.32 | 94.86
3.9 | 0.7 | 3.05 | 7.5 | 60.38 | 90.09 | 94.12 | 94.67
4 | 0.7 | 2.82 | 7 | 50.4 | 89.47 | 93.83 | 94.35
4.1 | 0.7 | 2.72 | 6.5 | 40.36 | 88.67 | 93.53 | 94.07
4.2 | 0.6 | 2.6 | 6.06 | 30.91 | 87.83 | 93.23 | 93.88
4.3 | 0.6 | 2.5 | 5.8 | 23.29 | 86.62 | 92.93 | 93.57
4.4 | 0.6 | 2.4 | 5.5 | 17.55 | 85.26 | 92.63 | 93.38
4.5 | 0.6 | 2.3 | 5.2 | 14.02 | 83.07 | 92.33 | 93.07
4.6 | 0.6 | 2.2 | 5 | 11.93 | 79.63 | 92.03 | 92.77
4.7 | 0.6 | 2.09 | 4.8 | 10.6 | 74.38 | 91.73 | 92.57
4.8 | 0.5 | 2.01 | 4.57 | 9.63 | 67.24 | 91.44 | 92.36
4.9 |  | 2.01 | 4.37 | 8.94 | 58.46 | 91.04 | 91.98
5 |  | 1.89 | 4.2 | 8.45 | 48.91 | 90.73 | 91.69
5.1 |  | 1.81 | 3.99 | 7.95 | 39.39 | 90.43 | 91.47
5.2 |  | 1.81 | 3.9 | 7.47 | 30.62 | 90.13 | 91.17
5.3 |  | 1.68 | 3.7 | 7.17 | 23.43 | 89.83 | 90.85
5.4 |  | 1.62 | 3.6 | 6.87 | 18.06 | 89.54 | 90.58
5.5 |  | 1.6 | 3.4 | 6.57 | 14.8 | 89.14 | 90.27
5.6 |  | 1.54 | 3.3 | 6.3 | 12.84 | 88.83 | 89.97
5.7 |  | 1.5 | 3.2 | 6 | 11.59 | 88.53 | 89.76
5.8 |  | 1.44 | 3.1 | 5.8 | 10.69 | 88.14 | 89.47
5.9 |  | 1.4 | 3 | 5.6 | 10.03 | 87.73 | 89.07
6 |  | 1.34 | 2.9 | 5.4 | 9.45 | 87.43 | 88.87
6.1 |  | 1.3 | 2.8 | 5.2 | 8.94 | 87.04 | 88.47
6.2 |  | 1.3 | 2.7 | 5 | 8.55 | 86.72 | 88.26
6.3 |  | 1.22 | 2.6 | 4.8 | 8.15 | 86.44 | 87.95
6.4 |  | 1.25 | 2.5 | 4.7 | 7.77 | 85.94 | 87.67
6.5 |  | 1.2 | 2.5 | 4.5 | 7.55 | 85.64 | 87.37
6.6 |  | 1.13 | 2.4 | 4.4 | 7.28 | 85.14 | 86.98
6.7 |  | 1.1 | 2.3 | 4.2 | 6.97 | 84.83 | 86.68
6.8 |  | 1.1 | 2.2 | 4.1 | 6.77 | 84.44 | 86.37
6.9 |  | 1.1 | 2.2 | 3.9 | 6.47 | 84.04 | 86.17
7 |  | 1.06 | 2.1 | 3.8 | 6.27 | 83.63 | 85.76
7.1 |  | 1 | 2 | 3.7 | 6.08 | 83.34 | 85.49
7.2 |  | 1 | 2 | 3.6 | 5.87 | 82.84 | 85.18
7.3 |  | 1 | 2 | 3.5 | 5.7 | 82.54 | 84.88
7.4 |  | 0.94 | 1.9 | 3.4 | 5.57 | 82.05 | 84.47
7.5 |  | 0.9 | 1.8 | 3.3 | 5.38 | 81.64 | 84.17
7.6 |  | 0.9 | 1.8 | 3.2 | 5.19 | 81.25 | 83.97
7.7 |  | 0.9 | 1.7 | 3.1 | 5.07 | 80.75 | 83.56
7.8 |  | 0.85 | 1.7 | 3 | 4.9 | 80.34 | 83.26
7.9 |  | 0.8 | 1.7 | 2.9 | 4.78 | 79.85 | 82.97
8 |  | 0.8 | 1.6 | 2.8 | 4.6 | 79.44 | 82.58
8.1 |  | 0.8 | 1.6 | 2.8 | 4.5 | 79.05 | 82.27
8.2 |  | 0.77 | 1.5 | 2.7 | 4.4 | 78.45 | 81.96
8.3 |  | 0.75 | 1.5 | 2.6 | 4.3 | 78.04 | 81.58
8.4 |  | 0.7 | 1.5 | 2.6 | 4.2 | 77.64 | 81.28
8.5 |  | 0.7 | 1.4 | 2.5 | 4 | 77.16 | 80.98
8.6 |  | 0.7 | 1.4 | 2.4 | 4 | 76.56 | 80.68
8.7 |  | 0.7 | 1.4 | 2.4 | 3.8 | 76.05 | 80.37
8.8 |  | 0.7 | 1.3 | 2.3 | 3.7 | 75.56 | 80.06
8.9 |  |  | 1.3 | 2.3 | 3.7 | 74.87 | 79.66
9 |  |  | 1.3 | 2.2 | 3.6 | 74.27 | 79.36
9.1 |  |  | 1.19 | 2.2 | 3.5 | 73.48 | 79.07
9.2 |  |  | 1.2 | 2.1 | 3.4 | 72.69 | 78.66
9.3 |  |  | 1.2 | 2 | 3.3 | 71.7 | 78.37
9.4 |  |  | 1.2 | 2 | 3.3 | 70.62 | 77.96
9.5 |  |  | 1.1 | 1.9 | 3.2 | 68.39 | 77.67
9.6 |  |  | 1.1 | 1.9 | 3.1 | 65.14 | 77.26
9.7 |  |  | 1.1 | 1.9 | 3 | 61.22 | 76.96
9.8 |  |  | 1.1 | 1.8 | 2.9 | 54.42 | 76.67
9.9 |  |  |  | 1.8 | 2.9 | 47.84 | 76.28
10 |  |  |  | 1.7 | 2.8 | 40.55 | 75.96
10.1 |  |  |  | 1.7 | 2.8 | 33.02 | 75.56
10.2 |  |  |  | 1.7 | 2.7 | 26.38 | 75.28
10.3 |  |  |  | 1.6 | 2.6 | 21.08 | 74.86
10.4 |  |  |  | 1.6 | 2.6 | 17.34 | 74.46
10.5 |  |  |  | 1.6 | 2.5 | 15.21 | 74.08
10.6 |  |  |  | 1.5 | 2.5 | 13.82 | 73.75
10.7 |  |  |  | 1.5 | 2.4 | 12.94 | 73.47
10.8 |  |  |  | 1.5 | 2.4 | 12.23 | 73.08
10.9 |  |  |  |  | 2.3 | 11.7 | 72.76
11 |  |  |  |  | 2.3 | 11.19 | 72.35
11.1 |  |  |  |  | 2.2 | 10.78 | 72.08
11.2 |  |  |  |  | 2.2 | 10.37 | 71.66
11.3 |  |  |  |  | 2.2 | 10.06 | 71.27
11.4 |  |  |  |  | 2.1 | 9.76 | 70.86
11.5 |  |  |  |  | 2.1 | 9.45 | 70.46
11.6 |  |  |  |  | 2.1 | 9.25 | 70.07
11.7 |  |  |  |  | 2 | 8.96 | 69.74
11.8 |  |  |  |  | 2 | 8.74 | 69.38
11.9 | [5 empty] | 8.46 | 68.96
12 | [5 empty] | 8.24 | 68.54
12.1 | [5 empty] | 8.04 | 68.18
12.2 | [5 empty] | 7.84 | 67.77
12.3 | [5 empty] | 7.63 | 67.25
12.4 | [5 empty] | 7.52 | 66.84
12.5 | [5 empty] | 7.34 | 66.46
12.6 | [5 empty] | 7.14 | 65.95
12.7 | [5 empty] | 6.93 | 65.5
12.8 | [5 empty] | 6.83 | 65.05
12.9 | [5 empty] | 6.63 | 64.44
13 | [5 empty] | 6.52 | 63.95
13.1 | [5 empty] | 6.34 | 63.4
13.2 | [5 empty] | 6.22 | 62.73
13.3 | [5 empty] | 6.05 | 62
13.4 | [5 empty] | 5.92 | 61.1
13.5 | [5 empty] | 5.83 | 59.8
13.6 | [5 empty] | 5.72 | 57.7
13.7 | [5 empty] | 5.62 | 54.4
13.8 | [5 empty] | 5.53 | 49.9
13.9 | [5 empty] | 5.35 | 44.4
14 | [5 empty] | 5.23 | 38.3
14.1 | [5 empty] | 5.11 | 32.1
14.2 | [5 empty] | 5.11 | 26.2
14.3 | [5 empty] | 4.96 | 21.3
14.4 | [5 empty] | 4.82 | 17.6
14.5 | [5 empty] | 4.72 | 15.2
14.6 | [5 empty] | 4.7 | 13.8
14.7 | [5 empty] | 4.63 | 12.9
14.8 | [5 empty] | 4.53 | 12.2
14.9 | [5 empty] | 4.43 | 11.6
15 | [5 empty] | 4.31 | 11.2
15.1 | [5 empty] | 4.3 | 10.8
15.2 | [5 empty] | 4.24 | 10.5
15.3 | [5 empty] | 4.12 | 10.1
15.4 | [5 empty] | 4.1 | 9.8
15.5 | [5 empty] | 4.03 | 9.6
15.6 | [5 empty] | 3.93 | 9.3
15.7 | [5 empty] | 3.82 | 9.1
15.8 | [5 empty] | 3.81 | 8.9
15.9 | [5 empty] | 3.73 | 8.6
16 | [5 empty] | 3.7 | 8.4
16.1 | [5 empty] | 3.64 | 8.2
16.2 | [5 empty] | 3.52 | 8
16.3 | [5 empty] | 3.51 | 7.9
16.4 | [5 empty] | 3.43 | 7.7
16.5 | [5 empty] | 3.4 | 7.5
16.6 | [5 empty] | 3.33 | 7.4
16.7 | [5 empty] | 3.3 | 7.2
16.8 | [5 empty] | 3.25 | 7
16.9 | [6 empty] | 6.9
17 | [6 empty] | 6.7
17.1 | [6 empty] | 6.6
17.2 | [6 empty] | 6.5
17.3 | [6 empty] | 6.4
17.4 | [6 empty] | 6.3
17.5 | [6 empty] | 6.1
17.6 | [6 empty] | 6
17.7 | [6 empty] | 5.9
17.8 | [6 empty] | 5.8
17.9 | [6 empty] | 5.7
18 | [6 empty] | 5.61
18.1 | [6 empty] | 5.5
18.2 | [6 empty] | 5.41
18.3 | [6 empty] | 5.3
18.4 | [6 empty] | 5.2
18.5 | [6 empty] | 5.09
18.6 | [6 empty] | 5.01
18.7 | [6 empty] | 4.92
18.8 | [6 empty] | 4.82
18.9 | [6 empty] | 4.72
19 | [6 empty] | 4.71
19.1 | [6 empty] | 4.59
19.2 | [6 empty] | 4.51
19.3 | [6 empty] | 4.42
19.4 | [6 empty] | 4.41
19.5 | [6 empty] | 4.3
19.6 | [6 empty] | 4.2
19.7 | [6 empty] | 4.21
19.8 | [6 empty] | 4.09
19.9 | [6 empty] | 4.02
20 | [6 empty] | 4.01
20.1 | [6 empty] | 3.89
20.2 | [6 empty] | 3.82
20.3 | [6 empty] | 3.81
20.4 | [6 empty] | 3.7
20.5 | [6 empty] | 3.71
20.6 | [6 empty] | 3.59
20.7 | [6 empty] | 3.52
20.8
20.9
21
21.1
21.2
21.3
21.4
21.5
21.6
21.7
21.8
21.9
22
22.1
22.2
22.3
22.4
22.5
22.6
22.7
22.8
22.9
23
23.1
23.2
23.3
23.4
23.5
23.6
23.7
23.8
23.9
24
24.1
24.2
24.3
24.4
24.5
24.6
24.7
24.8
24.9
25
25.1
25.2

## Open Field Profiles at 20cm
Lateral Profiles
Open field
SSD = 90 cm
Depth = 20 cm
Scanning direction: Crossline (i.e. along the leaf travel direction)
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 2x2 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 28x28 cm2
-25.2
-25.1
-25
-24.9
-24.8
-24.7
-24.6
-24.5
-24.4
-24.3
-24.2
-24.1
-24
-23.9
-23.8
-23.7
-23.6
-23.5
-23.4
-23.3
-23.2
-23.1
-23
-22.9
-22.8
-22.7
-22.6
-22.5
-22.4
-22.3 | [6 empty] | 5.94
-22.2 | [6 empty] | 5.99
-22.1 | [6 empty] | 6.1
-22 | [6 empty] | 6.2
-21.9 | [6 empty] | 6.21
-21.8 | [6 empty] | 6.31
-21.7 | [6 empty] | 6.41
-21.6 | [6 empty] | 6.51
-21.5 | [6 empty] | 6.61
-21.4 | [6 empty] | 6.71
-21.3 | [6 empty] | 6.79
-21.2 | [6 empty] | 6.92
-21.1 | [6 empty] | 7.01
-21 | [6 empty] | 7.09
-20.9 | [6 empty] | 7.22
-20.8 | [6 empty] | 7.31
-20.7 | [6 empty] | 7.41
-20.6 | [6 empty] | 7.51
-20.5 | [6 empty] | 7.6
-20.4 | [6 empty] | 7.8
-20.3 | [6 empty] | 7.92
-20.2 | [6 empty] | 8.01
-20.1 | [6 empty] | 8.1
-20 | [6 empty] | 8.29
-19.9 | [6 empty] | 8.4
-19.8 | [6 empty] | 8.51
-19.7 | [6 empty] | 8.69
-19.6 | [6 empty] | 8.81
-19.5 | [6 empty] | 8.99
-19.4 | [6 empty] | 9.1
-19.3 | [6 empty] | 9.21
-19.2 | [6 empty] | 9.39
-19.1 | [6 empty] | 9.51
-19 | [6 empty] | 9.71
-18.9 | [6 empty] | 9.8
-18.8 | [6 empty] | 10.01
-18.7 | [6 empty] | 10.19
-18.6 | [6 empty] | 10.31
-18.5 | [6 empty] | 10.51
-18.4 | [6 empty] | 10.71
-18.3 | [6 empty] | 10.91
-18.2 | [6 empty] | 10.99
-18.1 | [6 empty] | 11.3
-18 | [6 empty] | 11.41
-17.9 | [5 empty] | 5.59 | 11.72
-17.8 | [5 empty] | 5.66 | 11.91
-17.7 | [5 empty] | 5.69 | 12.11
-17.6 | [5 empty] | 5.77 | 12.31
-17.5 | [5 empty] | 5.87 | 12.49
-17.4 | [5 empty] | 5.97 | 12.79
-17.3 | [5 empty] | 6.07 | 12.99
-17.2 | [5 empty] | 6.17 | 13.32
-17.1 | [5 empty] | 6.27 | 13.49
-17 | [5 empty] | 6.37 | 13.79
-16.9 | [5 empty] | 6.46 | 14.09
-16.8 | [5 empty] | 6.57 | 14.39
-16.7 | [5 empty] | 6.66 | 14.68
-16.6 | [5 empty] | 6.77 | 15.08
-16.5 | [5 empty] | 6.86 | 15.38
-16.4 | [5 empty] | 6.97 | 15.78
-16.3 | [5 empty] | 7.06 | 16.28
-16.2 | [5 empty] | 7.17 | 16.78
-16.1 | [5 empty] | 7.26 | 17.48
-16 | [5 empty] | 7.45 | 18.48
-15.9 | [5 empty] | 7.57 | 19.78
-15.8 | [5 empty] | 7.66 | 21.88
-15.7 | [5 empty] | 7.75 | 24.97
-15.6 | [5 empty] | 7.93 | 28.97
-15.5 | [5 empty] | 8.05 | 33.57
-15.4 | [5 empty] | 8.23 | 38.66
-15.3 | [5 empty] | 8.36 | 43.55
-15.2 | [5 empty] | 8.54 | 48.15
-15.1 | [5 empty] | 8.65 | 51.95
-15 | [5 empty] | 8.83 | 54.94
-14.9 | [5 empty] | 8.95 | 56.84
-14.8 | [5 empty] | 9.13 | 58.24
-14.7 | [5 empty] | 9.25 | 59.14
-14.6 | [5 empty] | 9.41 | 60.04
-14.5 | [5 empty] | 9.64 | 60.54
-14.4 | [5 empty] | 9.75 | 61.24
-14.3 | [5 empty] | 9.92 | 61.74
-14.2 | [5 empty] | 10.13 | 62.24
-14.1 | [5 empty] | 10.32 | 62.73
-14 | [5 empty] | 10.54 | 63.23
-13.9 | [5 empty] | 10.64 | 63.63
-13.8 | [5 empty] | 10.89 | 64.13
-13.7 | [5 empty] | 11.13 | 64.53
-13.6 | [5 empty] | 11.32 | 64.93
-13.5 | [5 empty] | 11.52 | 65.43
-13.4 | [5 empty] | 11.71 | 65.83
-13.3 | [5 empty] | 12 | 66.23
-13.2 | [5 empty] | 12.22 | 66.63
-13.1 | [5 empty] | 12.49 | 67.03
-13 | [5 empty] | 12.72 | 67.43
-12.9 | [5 empty] | 12.98 | 67.83
-12.8 | [5 empty] | 13.3 | 68.23
-12.7 | [5 empty] | 13.5 | 68.53
-12.6 | [5 empty] | 13.86 | 69.03
-12.5 | [5 empty] | 14.18 | 69.43
-12.4 |  |  |  |  | 3.74 | 14.56 | 69.73
-12.3 |  |  |  |  | 3.8 | 14.89 | 70.13
-12.2 |  |  |  |  | 3.81 | 15.18 | 70.53
-12.1 |  |  |  |  | 3.91 | 15.54 | 70.93
-12 |  |  |  |  | 4.01 | 16.02 | 71.33
-11.9 |  |  |  |  | 4.1 | 16.52 | 71.53
-11.8 |  |  |  |  | 4.09 | 17.06 | 71.83
-11.7 |  |  |  |  | 4.22 | 17.88 | 72.32
-11.6 |  |  |  |  | 4.34 | 18.84 | 72.62
-11.5 |  |  |  |  | 4.31 | 20.13 | 73.02
-11.4 |  |  |  |  | 4.4 | 22.27 | 73.32
-11.3 |  |  |  | 2.83 | 4.5 | 25.25 | 73.72
-11.2 |  |  |  | 2.9 | 4.5 | 29.97 | 74.12
-11.1 |  |  |  | 2.91 | 4.73 | 35.2 | 74.42
-11 |  |  |  | 3.01 | 4.71 | 41.15 | 74.72
-10.9 |  |  |  | 3.1 | 4.81 | 47.05 | 75.02
-10.8 |  |  |  | 3.11 | 4.91 | 53.39 | 75.42
-10.7 |  |  |  | 3.2 | 5.01 | 58.58 | 75.72
-10.6 |  |  |  | 3.21 | 5.11 | 62.56 | 76.12
-10.5 |  |  |  | 3.29 | 5.21 | 65.61 | 76.42
-10.4 |  |  |  | 3.43 | 5.31 | 67.59 | 76.82
-10.3 |  |  |  | 3.41 | 5.41 | 69.05 | 77.12
-10.2 |  |  | 2.15 | 3.5 | 5.5 | 70.07 | 77.52
-10.1 |  |  | 2.2 | 3.51 | 5.73 | 71.07 | 77.82
-10 |  |  | 2.2 | 3.61 | 5.7 | 71.49 | 78.12
-9.9 |  |  | 2.28 | 3.7 | 5.94 | 72.39 | 78.42
-9.8 |  |  | 2.29 | 3.78 | 6.01 | 72.79 | 78.82
-9.7 |  |  | 2.28 | 3.97 | 6.11 | 73.41 | 79.02
-9.6 |  |  | 2.4 | 3.91 | 6.21 | 74.04 | 79.42
-9.5 |  |  | 2.4 | 4 | 6.4 | 74.39 | 79.82
-9.4 |  |  | 2.48 | 4.1 | 6.41 | 75.13 | 80.12
-9.3 |  |  | 2.48 | 4.2 | 6.6 | 75.34 | 80.52
-9.2 |  |  | 2.6 | 4.3 | 6.84 | 76.11 | 80.72
-9.1 |  | 1.36 | 2.58 | 4.4 | 6.91 | 76.46 | 81.12
-9 |  | 1.4 | 2.7 | 4.49 | 7.1 | 76.8 | 81.42
-8.9 |  | 1.4 | 2.68 | 4.68 | 7.19 | 77.33 | 81.72
-8.8 |  | 1.5 | 2.8 | 4.71 | 7.48 | 77.94 | 82.11
-8.7 |  | 1.36 | 2.88 | 4.79 | 7.51 | 78.26 | 82.41
-8.6 |  | 1.5 | 2.9 | 4.98 | 7.69 | 78.53 | 82.71
-8.5 |  | 1.59 | 2.98 | 5.01 | 7.96 | 79.1 | 83.01
-8.4 |  | 1.5 | 2.98 | 5.1 | 8.17 | 79.53 | 83.31
-8.3 |  | 1.5 | 3.1 | 5.2 | 8.38 | 80.13 | 83.61
-8.2 |  | 1.6 | 3.18 | 5.29 | 8.42 | 80.26 | 83.81
-8.1 |  | 1.6 | 3.18 | 5.58 | 8.7 | 80.8 | 84.18
-8 |  | 1.7 | 3.3 | 5.61 | 8.9 | 81.35 | 84.37
-7.9 |  | 1.7 | 3.38 | 5.7 | 9.18 | 81.53 | 84.8
-7.8 |  | 1.79 | 3.48 | 5.81 | 9.3 | 82.02 | 85.08
-7.7 |  | 1.81 | 3.58 | 6 | 9.51 | 82.43 | 85.38
-7.6 |  | 1.79 | 3.68 | 6.2 | 9.8 | 82.95 | 85.75
-7.5 |  | 1.9 | 3.69 | 6.31 | 10.1 | 83.13 | 86.09
-7.4 |  | 1.9 | 3.78 | 6.5 | 10.3 | 83.75 | 86.29
-7.3 |  | 1.9 | 3.88 | 6.6 | 10.6 | 84.05 | 86.49
-7.2 |  | 2 | 3.96 | 6.81 | 10.9 | 84.25 | 86.77
-7.1 |  | 2 | 4.09 | 7 | 11.3 | 84.85 | 87.18
-7 |  | 2.1 | 4.2 | 7.2 | 11.5 | 85.02 | 87.47
-6.9 |  | 2.1 | 4.27 | 7.4 | 11.8 | 85.51 | 87.68
-6.8 |  | 2.2 | 4.48 | 7.6 | 12.3 | 85.92 | 88.08
-6.7 |  | 2.2 | 4.47 | 7.8 | 12.7 | 86.14 | 88.28
-6.6 |  | 2.2 | 4.59 | 8 | 13.1 | 86.63 | 88.59
-6.5 |  | 2.29 | 4.77 | 8.2 | 13.6 | 86.86 | 88.79
-6.4 |  | 2.4 | 4.86 | 8.5 | 14.2 | 87.19 | 89.06
-6.3 |  | 2.49 | 4.94 | 8.7 | 15 | 87.53 | 89.38
-6.2 |  | 2.39 | 5.17 | 9 | 16 | 87.89 | 89.77
-6.1 |  | 2.49 | 5.24 | 9.2 | 17.3 | 88.24 | 89.98
-6 |  | 2.69 | 5.47 | 9.5 | 19.4 | 88.52 | 90.29
-5.9 |  | 2.69 | 5.54 | 9.8 | 22.8 | 88.89 | 90.56
-5.8 |  | 2.69 | 5.78 | 10.2 | 27.7 | 89.24 | 90.89
-5.7 |  | 2.89 | 5.93 | 10.5 | 34.3 | 89.53 | 91.08
-5.6 |  | 2.89 | 6.04 | 10.9 | 42.1 | 89.82 | 91.28
-5.5 |  | 2.89 | 6.25 | 11.4 | 50.5 | 90.2 | 91.59
-5.4 |  | 2.99 | 6.45 | 12 | 58.9 | 90.47 | 91.87
-5.3 |  | 3.19 | 6.66 | 12.5 | 66.7 | 90.72 | 92.08
-5.2 |  | 3.19 | 6.92 | 13.4 | 73.2 | 91.02 | 92.27
-5.1 |  | 3.29 | 7.17 | 14.4 | 78 | 91.39 | 92.58
-5 | 0.9 | 3.49 | 7.31 | 15.9 | 81.3 | 91.74 | 92.89
-4.9 | 0.89 | 3.49 | 7.62 | 18.1 | 83.6 | 91.97 | 93.09
-4.8 | 1.02 | 3.69 | 7.92 | 21.8 | 85.1 | 92.21 | 93.37
-4.7 | 0.93 | 3.79 | 8.22 | 27.2 | 86.3 | 92.6 | 93.58
-4.6 | 1 | 3.89 | 8.5 | 34.3 | 87.3 | 92.89 | 93.78
-4.5 | 1.03 | 3.99 | 8.91 | 42.7 | 88 | 93.04 | 94.07
-4.4 | 1.09 | 4.09 | 9.32 | 51.7 | 88.7 | 93.39 | 94.28
-4.3 | 1.1 | 4.19 | 9.85 | 60.9 | 89.3 | 93.66 | 94.48
-4.2 | 1.13 | 4.49 | 10.37 | 69.2 | 89.9 | 93.94 | 94.68
-4.1 | 1.19 | 4.49 | 11.17 | 76.1 | 90.4 | 94.09 | 94.96
-4 | 1.19 | 4.79 | 12.15 | 81.2 | 90.9 | 94.38 | 95.18
-3.9 | 1.32 | 4.89 | 13.61 | 84.7 | 91.3 | 94.67 | 95.38
-3.8 | 1.22 | 5.09 | 15.89 | 86.9 | 91.9 | 94.92 | 95.56
-3.7 | 1.4 | 5.39 | 19.53 | 88.5 | 92.4 | 95.16 | 95.78
-3.6 | 1.39 | 5.69 | 25.03 | 89.6 | 92.8 | 95.42 | 96.04
-3.5 | 1.51 | 5.79 | 32.33 | 90.5 | 93.2 | 95.66 | 96.26
-3.4 | 1.51 | 6.19 | 41.37 | 91.4 | 93.6 | 95.89 | 96.4
-3.3 | 1.6 | 6.39 | 50.97 | 92 | 94 | 96.02 | 96.62
-3.2 | 1.69 | 6.79 | 60.93 | 92.6 | 94.4 | 96.3 | 96.8
-3.1 | 1.71 | 7.08 | 69.41 | 93 | 94.6 | 96.42 | 96.97
-3 | 1.79 | 7.68 | 77.67 | 93.6 | 95 | 96.7 | 97.05
-2.9 | 1.91 | 8.38 | 83.04 | 94 | 95.5 | 96.81 | 97.26
-2.8 | 1.99 | 8.92 | 87.23 | 94.5 | 95.7 | 97.16 | 97.46
-2.7 | 2.01 | 9.94 | 89.91 | 94.9 | 96.1 | 97.27 | 97.67
-2.6 | 2.19 | 12.07 | 91.69 | 95.35 | 96.3 | 97.46 | 97.75
-2.5 | 2.29 | 15.38 | 92.82 | 95.6 | 96.6 | 97.65 | 97.97
-2.4 | 2.49 | 23.89 | 93.84 | 96 | 96.9 | 97.89 | 98.06
-2.3 | 2.59 | 37.82 | 94.62 | 96.2 | 97.2 | 98 | 98.15
-2.2 | 2.79 | 52.6 | 95.13 | 96.7 | 97.5 | 98.15 | 98.35
-2.1 | 3.09 | 68.75 | 95.68 | 96.9 | 97.7 | 98.29 | 98.49
-2 | 3.29 | 80.45 | 96.2 | 97.2 | 98 | 98.45 | 98.6
-1.9 | 3.69 | 87.69 | 96.56 | 97.5 | 98.1 | 98.57 | 98.7
-1.8 | 4.09 | 91.96 | 97 | 97.7 | 98.3 | 98.72 | 98.86
-1.7 | 4.88 | 93.96 | 97.28 | 98 | 98.5 | 98.82 | 99.08
-1.6 | 5.88 | 95.13 | 97.63 | 98.26 | 98.7 | 99.02 | 99.04
-1.5 | 7.97 | 96.15 | 97.92 | 98.4 | 98.9 | 99.07 | 99.19
-1.4 | 12.26 | 96.75 | 98.17 | 98.54 | 99 | 99.21 | 99.37
-1.3 | 21.33 | 97.14 | 98.37 | 98.8 | 99.2 | 99.29 | 99.4
-1.2 | 36.18 | 97.89 | 98.53 | 99 | 99.3 | 99.39 | 99.4
-1.1 | 54.32 | 97.9 | 98.78 | 99.2 | 99.4 | 99.43 | 99.45
-1 | 71.66 | 98.73 | 99.02 | 99.25 | 99.5 | 99.49 | 99.6
-0.9 | 83.54 | 98.43 | 99.19 | 99.4 | 99.62 | 99.59 | 99.6
-0.8 | 91.16 | 98.39 | 99.26 | 99.5 | 99.8 | 99.69 | 99.65
-0.7 | 95.35 | 99.06 | 99.49 | 99.6 | 99.8 | 99.73 | 99.8
-0.6 | 97.2 | 98.8 | 99.51 | 99.7 | 99.9 | 99.73 | 99.85
-0.5 | 98.69 | 99.45 | 99.63 | 99.8 | 99.9 | 99.79 | 99.9
-0.4 | 99.3 | 99.58 | 99.78 | 99.9 | 99.99 | 99.83 | 99.96
-0.3 | 99.42 | 99.63 | 100 | 99.9 | 99.9 | 99.77 | 100
-0.2 | 99.74 | 99.69 | 99.81 | 100 | 100 | 99.67 | 99.89
-0.1 | 100 | 100 | 99.78 | 100 | 100 | 100 | 99.9
0 | 99.67 | 99.78 | 99.82 | 100 | 100 | 99.83 | 99.9
0.1 | 100 | 100 | 99.78 | 100 | 100 | 100 | 99.9
0.2 | 99.74 | 99.69 | 99.81 | 100 | 100 | 99.68 | 99.89
0.3 | 99.42 | 99.63 | 100 | 99.9 | 99.95 | 99.79 | 100
0.4 | 99.3 | 99.58 | 99.78 | 99.9 | 100 | 99.83 | 99.95
0.5 | 98.69 | 99.45 | 99.63 | 99.8 | 99.9 | 99.78 | 99.9
0.6 | 97.2 | 98.82 | 99.51 | 99.7 | 99.9 | 99.74 | 99.85
0.7 | 95.35 | 99.07 | 99.49 | 99.6 | 99.8 | 99.73 | 99.8
0.8 | 91.16 | 98.39 | 99.26 | 99.5 | 99.8 | 99.68 | 99.65
0.9 | 83.55 | 98.43 | 99.19 | 99.4 | 99.6 | 99.58 | 99.6
1 | 70.48 | 98.73 | 99.02 | 99.25 | 99.5 | 99.49 | 99.6
1.1 | 54.32 | 97.9 | 98.78 | 99.2 | 99.4 | 99.43 | 99.45
1.2 | 36.18 | 97.89 | 98.53 | 99 | 99.3 | 99.39 | 99.4
1.3 | 21.33 | 97.14 | 98.37 | 98.8 | 99.2 | 99.29 | 99.4
1.4 | 12.26 | 96.75 | 98.17 | 98.54 | 99 | 99.21 | 99.37
1.5 | 7.97 | 96.15 | 97.92 | 98.4 | 98.9 | 99.07 | 99.19
1.6 | 5.88 | 95.13 | 97.63 | 98.26 | 98.7 | 99.02 | 99.04
1.7 | 4.88 | 93.96 | 97.28 | 98 | 98.5 | 98.8 | 99.08
1.8 | 4.09 | 91.96 | 97 | 97.7 | 98.3 | 98.72 | 98.86
1.9 | 3.69 | 87.69 | 96.56 | 97.5 | 98.1 | 98.57 | 98.7
2 | 3.29 | 80.45 | 96.2 | 97.2 | 98 | 98.45 | 98.6
2.1 | 3.09 | 68.75 | 95.68 | 96.9 | 97.7 | 98.29 | 98.49
2.2 | 2.79 | 52.61 | 95.13 | 96.65 | 97.5 | 98.15 | 98.35
2.3 | 2.59 | 36.36 | 94.62 | 96.2 | 97.2 | 98 | 98.15
2.4 | 2.49 | 23.39 | 93.84 | 96 | 96.9 | 97.89 | 98.06
2.5 | 2.29 | 15.4 | 92.82 | 95.6 | 96.6 | 97.64 | 97.97
2.6 | 2.19 | 11.82 | 91.69 | 95.35 | 96.3 | 97.44 | 97.75
2.7 | 2.01 | 9.95 | 89.91 | 94.9 | 96.1 | 97.27 | 97.67
2.8 | 1.99 | 8.92 | 87.24 | 94.5 | 95.7 | 97.16 | 97.46
2.9 | 1.91 | 8.38 | 83.21 | 94 | 95.5 | 96.81 | 97.26
3 | 1.79 | 7.62 | 77.15 | 93.6 | 95 | 96.7 | 97.05
3.1 | 1.71 | 7.08 | 69.62 | 93 | 94.6 | 96.42 | 96.97
3.2 | 1.69 | 6.72 | 60.08 | 92.6 | 94.4 | 96.3 | 96.8
3.3 | 1.6 | 6.32 | 50.93 | 92 | 94 | 96.02 | 96.62
3.4 | 1.51 | 6.19 | 41.37 | 91.4 | 93.6 | 95.89 | 96.4
3.5 | 1.51 | 5.79 | 32.33 | 90.5 | 93.2 | 95.66 | 96.26
3.6 | 1.39 | 5.69 | 25.03 | 89.6 | 92.8 | 95.42 | 96.04
3.7 | 1.4 | 5.39 | 19.53 | 88.5 | 92.4 | 95.16 | 95.78
3.8 | 1.22 | 5.09 | 15.89 | 86.9 | 91.9 | 94.92 | 95.56
3.9 | 1.32 | 4.89 | 13.61 | 84.7 | 91.3 | 94.67 | 95.38
4 | 1.19 | 4.79 | 12.15 | 81.2 | 90.9 | 94.38 | 95.18
4.1 | 1.19 | 4.49 | 11.17 | 76.1 | 90.4 | 94.09 | 94.96
4.2 | 1.13 | 4.5 | 10.37 | 69.2 | 89.9 | 93.94 | 94.68
4.3 | 1.1 | 4.19 | 9.85 | 60.9 | 89.3 | 93.66 | 94.48
4.4 | 1.09 | 4.09 | 9.32 | 51.7 | 88.7 | 93.39 | 94.28
4.5 | 1.03 | 3.99 | 8.91 | 42.7 | 88 | 93.04 | 94.07
4.6 | 1 | 3.89 | 8.5 | 34.3 | 87.3 | 92.89 | 93.78
4.7 | 0.93 | 3.79 | 8.22 | 27.2 | 86.3 | 92.6 | 93.58
4.8 | 1.02 | 3.69 | 7.92 | 21.8 | 85.1 | 92.21 | 93.37
4.9 | 0.89 | 3.49 | 7.62 | 18.1 | 83.6 | 91.97 | 93.09
5 |  | 3.49 | 7.31 | 15.9 | 81.3 | 91.74 | 92.89
5.1 |  | 3.29 | 7.17 | 14.4 | 78 | 91.39 | 92.58
5.2 |  | 3.19 | 6.92 | 13.4 | 73.2 | 91.02 | 92.27
5.3 |  | 3.19 | 6.66 | 12.5 | 66.7 | 90.72 | 92.08
5.4 |  | 2.99 | 6.45 | 12 | 58.9 | 90.47 | 91.87
5.5 |  | 2.89 | 6.25 | 11.4 | 50.5 | 90.2 | 91.59
5.6 |  | 2.89 | 6.04 | 10.9 | 42.1 | 89.82 | 91.28
5.7 |  | 2.89 | 5.93 | 10.5 | 34.3 | 89.53 | 91.08
5.8 |  | 2.69 | 5.78 | 10.2 | 27.7 | 89.24 | 90.89
5.9 |  | 2.69 | 5.54 | 9.8 | 22.8 | 88.89 | 90.56
6 |  | 2.69 | 5.47 | 9.5 | 19.4 | 88.52 | 90.29
6.1 |  | 2.49 | 5.24 | 9.2 | 17.3 | 88.24 | 89.98
6.2 |  | 2.39 | 5.17 | 9 | 16 | 87.89 | 89.77
6.3 |  | 2.49 | 4.94 | 8.7 | 15 | 87.53 | 89.38
6.4 |  | 2.4 | 4.86 | 8.5 | 14.2 | 87.19 | 89.06
6.5 |  | 2.29 | 4.77 | 8.2 | 13.6 | 86.86 | 88.79
6.6 |  | 2.2 | 4.59 | 8 | 13.1 | 86.63 | 88.59
6.7 |  | 2.2 | 4.47 | 7.8 | 12.7 | 86.14 | 88.28
6.8 |  | 2.2 | 4.48 | 7.6 | 12.3 | 85.92 | 88.08
6.9 |  | 2.1 | 4.27 | 7.4 | 11.8 | 85.51 | 87.68
7 |  | 2.1 | 4.2 | 7.2 | 11.5 | 85.02 | 87.47
7.1 |  | 2 | 4.09 | 7 | 11.3 | 84.85 | 87.18
7.2 |  | 2 | 3.97 | 6.81 | 10.9 | 84.25 | 86.77
7.3 |  | 1.9 | 3.88 | 6.6 | 10.6 | 84.05 | 86.49
7.4 |  | 1.9 | 3.78 | 6.5 | 10.3 | 83.75 | 86.29
7.5 |  | 1.9 | 3.69 | 6.31 | 10.1 | 83.13 | 86.09
7.6 |  | 1.8 | 3.68 | 6.2 | 9.8 | 82.95 | 85.75
7.7 |  | 1.81 | 3.58 | 6 | 9.5 | 82.43 | 85.38
7.8 |  | 1.79 | 3.48 | 5.81 | 9.3 | 82.02 | 85.08
7.9 |  | 1.7 | 3.38 | 5.7 | 9.18 | 81.53 | 84.8
8 |  | 1.7 | 3.3 | 5.61 | 8.9 | 81.35 | 84.37
8.1 |  | 1.6 | 3.18 | 5.58 | 8.7 | 80.8 | 84.18
8.2 |  | 1.6 | 3.18 | 5.29 | 8.41 | 80.26 | 83.79
8.3 |  | 1.5 | 3.1 | 5.2 | 8.3 | 80.13 | 83.58
8.4 |  | 1.5 | 2.98 | 5.1 | 8.1 | 79.53 | 83.3
8.5 |  | 1.6 | 2.98 | 5.01 | 7.97 | 79.1 | 83.01
8.6 |  | 1.5 | 2.9 | 4.98 | 7.69 | 78.53 | 82.71
8.7 |  | 1.36 | 2.88 | 4.79 | 7.51 | 78.26 | 82.41
8.8 |  | 1.5 | 2.8 | 4.71 | 7.48 | 77.94 | 82.11
8.9 |  | 1.4 | 2.69 | 4.68 | 7.19 | 77.33 | 81.72
9 |  | 1.4 | 2.7 | 4.49 | 7.1 | 76.8 | 81.42
9.1 |  |  | 2.59 | 4.4 | 6.91 | 76.46 | 81.12
9.2 |  |  | 2.6 | 4.3 | 6.84 | 76.11 | 80.72
9.3 |  |  | 2.48 | 4.2 | 6.6 | 75.34 | 80.52
9.4 |  |  | 2.48 | 4.1 | 6.41 | 75.13 | 80.12
9.5 |  |  | 2.4 | 4 | 6.4 | 74.39 | 79.82
9.6 |  |  | 2.4 | 3.91 | 6.21 | 74.04 | 79.42
9.7 |  |  | 2.28 | 3.97 | 6.11 | 73.41 | 79.02
9.8 |  |  | 2.29 | 3.79 | 6.01 | 72.79 | 78.82
9.9 |  |  | 2.28 | 3.7 | 5.94 | 72.39 | 78.42
10 |  |  | 2.2 | 3.6 | 5.7 | 71.49 | 78.12
10.1 |  |  | 2.2 | 3.51 | 5.73 | 71.07 | 77.82
10.2 |  |  |  | 3.5 | 5.5 | 70.07 | 77.52
10.3 |  |  |  | 3.41 | 5.41 | 69.05 | 77.12
10.4 |  |  |  | 3.43 | 5.31 | 67.59 | 76.82
10.5 |  |  |  | 3.29 | 5.21 | 65.61 | 76.42
10.6 |  |  |  | 3.21 | 5.11 | 62.56 | 76.12
10.7 |  |  |  | 3.2 | 5.01 | 58.58 | 75.72
10.8 |  |  |  | 3.11 | 4.91 | 53.39 | 75.42
10.9 |  |  |  | 3.1 | 4.81 | 47.05 | 75.02
11 |  |  |  | 3.01 | 4.71 | 41.15 | 74.72
11.1 |  |  |  | 2.91 | 4.73 | 35.2 | 74.42
11.2 |  |  |  | 2.9 | 4.5 | 29.97 | 74.12
11.3 |  |  |  |  | 4.5 | 25.25 | 73.72
11.4 |  |  |  |  | 4.4 | 22.27 | 73.32
11.5 |  |  |  |  | 4.31 | 20.13 | 73.02
11.6 |  |  |  |  | 4.34 | 18.84 | 72.62
11.7 |  |  |  |  | 4.22 | 17.88 | 72.32
11.8 |  |  |  |  | 4.09 | 17.06 | 71.83
11.9 |  |  |  |  | 4.1 | 16.52 | 71.53
12 |  |  |  |  | 4.01 | 16.02 | 71.33
12.1 |  |  |  |  | 3.91 | 15.54 | 70.93
12.2 |  |  |  |  | 3.81 | 15.18 | 70.53
12.3 |  |  |  |  | 3.8 | 14.89 | 70.13
12.4 | [5 empty] | 14.56 | 69.73
12.5 | [5 empty] | 14.18 | 69.43
12.6 | [5 empty] | 13.86 | 69.03
12.7 | [5 empty] | 13.5 | 68.53
12.8 | [5 empty] | 13.3 | 68.23
12.9 | [5 empty] | 12.98 | 67.83
13 | [5 empty] | 12.72 | 67.43
13.1 | [5 empty] | 12.49 | 67.03
13.2 | [5 empty] | 12.22 | 66.63
13.3 | [5 empty] | 12 | 66.23
13.4 | [5 empty] | 11.71 | 65.83
13.5 | [5 empty] | 11.52 | 65.43
13.6 | [5 empty] | 11.32 | 64.93
13.7 | [5 empty] | 11.13 | 64.53
13.8 | [5 empty] | 10.89 | 64.13
13.9 | [5 empty] | 10.64 | 63.63
14 | [5 empty] | 10.54 | 63.23
14.1 | [5 empty] | 10.32 | 62.73
14.2 | [5 empty] | 10.13 | 62.24
14.3 | [5 empty] | 9.92 | 61.74
14.4 | [5 empty] | 9.75 | 61.24
14.5 | [5 empty] | 9.64 | 60.54
14.6 | [5 empty] | 9.41 | 60.04
14.7 | [5 empty] | 9.25 | 59.14
14.8 | [5 empty] | 9.13 | 58.24
14.9 | [5 empty] | 8.95 | 56.84
15 | [5 empty] | 8.83 | 54.94
15.1 | [5 empty] | 8.65 | 51.95
15.2 | [5 empty] | 8.54 | 48.15
15.3 | [5 empty] | 8.36 | 43.55
15.4 | [5 empty] | 8.23 | 38.66
15.5 | [5 empty] | 8.05 | 33.57
15.6 | [5 empty] | 7.93 | 28.97
15.7 | [5 empty] | 7.75 | 24.97
15.8 | [5 empty] | 7.66 | 21.88
15.9 | [5 empty] | 7.57 | 19.78
16 | [5 empty] | 7.45 | 18.48
16.1 | [5 empty] | 7.26 | 17.48
16.2 | [5 empty] | 7.17 | 16.78
16.3 | [5 empty] | 7.06 | 16.28
16.4 | [5 empty] | 6.97 | 15.78
16.5 | [5 empty] | 6.86 | 15.38
16.6 | [5 empty] | 6.77 | 15.08
16.7 | [5 empty] | 6.66 | 14.68
16.8 | [5 empty] | 6.57 | 14.39
16.9 | [5 empty] | 6.46 | 14.09
17 | [5 empty] | 6.37 | 13.79
17.1 | [5 empty] | 6.27 | 13.49
17.2 | [5 empty] | 6.17 | 13.32
17.3 | [5 empty] | 6.07 | 12.99
17.4 | [5 empty] | 5.97 | 12.79
17.5 | [5 empty] | 5.87 | 12.49
17.6 | [5 empty] | 5.77 | 12.31
17.7 | [5 empty] | 5.69 | 12.11
17.8 | [5 empty] | 5.66 | 11.91
17.9 | [6 empty] | 11.72
18 | [6 empty] | 11.41
18.1 | [6 empty] | 11.3
18.2 | [6 empty] | 10.99
18.3 | [6 empty] | 10.91
18.4 | [6 empty] | 10.71
18.5 | [6 empty] | 10.51
18.6 | [6 empty] | 10.31
18.7 | [6 empty] | 10.19
18.8 | [6 empty] | 10.01
18.9 | [6 empty] | 9.8
19 | [6 empty] | 9.71
19.1 | [6 empty] | 9.51
19.2 | [6 empty] | 9.39
19.3 | [6 empty] | 9.21
19.4 | [6 empty] | 9.1
19.5 | [6 empty] | 8.99
19.6 | [6 empty] | 8.81
19.7 | [6 empty] | 8.69
19.8 | [6 empty] | 8.51
19.9 | [6 empty] | 8.4
20 | [6 empty] | 8.29
20.1 | [6 empty] | 8.1
20.2 | [6 empty] | 8.01
20.3 | [6 empty] | 7.92
20.4 | [6 empty] | 7.8
20.5 | [6 empty] | 7.6
20.6 | [6 empty] | 7.51
20.7 | [6 empty] | 7.41
20.8 | [6 empty] | 7.31
20.9 | [6 empty] | 7.22
21 | [6 empty] | 7.09
21.1 | [6 empty] | 7.01
21.2 | [6 empty] | 6.92
21.3 | [6 empty] | 6.79
21.4 | [6 empty] | 6.71
21.5 | [6 empty] | 6.61
21.6 | [6 empty] | 6.51
21.7 | [6 empty] | 6.41
21.8 | [6 empty] | 6.31
21.9 | [6 empty] | 6.21
22 | [6 empty] | 6.2
22.1 | [6 empty] | 6.1
22.2 | [6 empty] | 5.99
22.3
22.4
22.5
22.6
22.7
22.8
22.9
23
23.1
23.2
23.3
23.4
23.5
23.6
23.7
23.8
23.9
24
24.1
24.2
24.3
24.4
24.5
24.6
24.7
24.8
24.9
25
25.1
25.2

## Open Field Profiles at 30cm
Lateral Profiles
Open field
SSD = 90 cm
Depth = 30 cm
Scanning direction: Crossline (i.e. along the leaf travel direction)
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 2x2 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 28x28 cm2
-25.2
-25.1
-25
-24.9
-24.8
-24.7
-24.6
-24.5
-24.4
-24.3
-24.2
-24.1
-24
-23.9
-23.8
-23.7
-23.6
-23.5
-23.4
-23.3
-23.2
-23.1
-23
-22.9
-22.8
-22.7 | [6 empty] | 9.18
-22.6 | [6 empty] | 9.28
-22.5 | [6 empty] | 9.48
-22.4 | [6 empty] | 9.59
-22.3 | [6 empty] | 9.69
-22.2 | [6 empty] | 9.79
-22.1 | [6 empty] | 9.88
-22 | [6 empty] | 10.08
-21.9 | [6 empty] | 10.19
-21.8 | [6 empty] | 10.28
-21.7 | [6 empty] | 10.48
-21.6 | [6 empty] | 10.59
-21.5 | [6 empty] | 10.78
-21.4 | [6 empty] | 10.89
-21.3 | [6 empty] | 10.99
-21.2 | [6 empty] | 11.08
-21.1 | [6 empty] | 11.38
-21 | [6 empty] | 11.49
-20.9 | [6 empty] | 11.68
-20.8 | [6 empty] | 11.79
-20.7 | [6 empty] | 11.89
-20.6 | [6 empty] | 12.07
-20.5 | [6 empty] | 12.38
-20.4 | [6 empty] | 12.49
-20.3 | [6 empty] | 12.59
-20.2 | [6 empty] | 12.77
-20.1 | [6 empty] | 13.08
-20 | [6 empty] | 13.19
-19.9 | [6 empty] | 13.37
-19.8 | [6 empty] | 13.68
-19.7 | [6 empty] | 13.79
-19.6 | [6 empty] | 13.98
-19.5 | [6 empty] | 14.18
-19.4 | [6 empty] | 14.38
-19.3 | [6 empty] | 14.57
-19.2 | [6 empty] | 14.87
-19.1 | [6 empty] | 15.08
-19 | [6 empty] | 15.28
-18.9 | [5 empty] | 7.6 | 15.55
-18.8 | [5 empty] | 7.8 | 15.76
-18.7 | [5 empty] | 7.88 | 16.03
-18.6 | [5 empty] | 7.96 | 16.35
-18.5 | [5 empty] | 8.09 | 16.56
-18.4 | [5 empty] | 8.16 | 16.84
-18.3 | [5 empty] | 8.29 | 17.14
-18.2 | [5 empty] | 8.38 | 17.44
-18.1 | [5 empty] | 8.54 | 17.73
-18 | [5 empty] | 8.69 | 18.12
-17.9 | [5 empty] | 8.77 | 18.44
-17.8 | [5 empty] | 8.88 | 18.81
-17.7 | [5 empty] | 8.97 | 19.29
-17.6 | [5 empty] | 9.16 | 19.88
-17.5 | [5 empty] | 9.29 | 20.53
-17.4 | [5 empty] | 9.37 | 21.56
-17.3 | [5 empty] | 9.46 | 22.87
-17.2 | [5 empty] | 9.69 | 24.9
-17.1 | [5 empty] | 9.66 | 27.67
-17 | [5 empty] | 10.02 | 31.11
-16.9 | [5 empty] | 10.06 | 35.34
-16.8 | [5 empty] | 10.26 | 39.46
-16.7 | [5 empty] | 10.36 | 43.84
-16.6 | [5 empty] | 10.58 | 47.8
-16.5 | [5 empty] | 10.66 | 51.03
-16.4 | [5 empty] | 10.85 | 53.66
-16.3 | [5 empty] | 11.06 | 55.51
-16.2 | [5 empty] | 11.18 | 56.78
-16.1 | [5 empty] | 11.27 | 57.83
-16 | [5 empty] | 11.45 | 58.48
-15.9 | [5 empty] | 11.66 | 59.16
-15.8 | [5 empty] | 11.85 | 59.8
-15.7 | [5 empty] | 12.05 | 60.2
-15.6 | [5 empty] | 12.26 | 60.8
-15.5 | [5 empty] | 12.39 | 61.14
-15.4 | [5 empty] | 12.65 | 61.6
-15.3 | [5 empty] | 12.87 | 62.01
-15.2 | [5 empty] | 13.03 | 62.52
-15.1 | [5 empty] | 13.28 | 62.82
-15 | [5 empty] | 13.43 | 63.38
-14.9 | [5 empty] | 13.68 | 63.84
-14.8 | [5 empty] | 13.86 | 64.13
-14.7 | [5 empty] | 13.87 | 64.52
-14.6 | [5 empty] | 14.31 | 64.92
-14.5 | [5 empty] | 14.47 | 65.32
-14.4 | [5 empty] | 14.75 | 65.72
-14.3 | [5 empty] | 14.85 | 66.12
-14.2 | [5 empty] | 15.31 | 66.51
-14.1 | [5 empty] | 15.55 | 67.02
-14 | [5 empty] | 15.83 | 67.27
-13.9 | [5 empty] | 15.96 | 67.53
-13.8 | [5 empty] | 16.38 | 67.92
-13.7 | [5 empty] | 16.75 | 68.33
-13.6 | [5 empty] | 17 | 68.64
-13.5 | [5 empty] | 17.37 | 69.01
-13.4 | [5 empty] | 17.6 | 69.43
-13.3 | [5 empty] | 18.03 | 69.75
-13.2 | [5 empty] | 18.32 | 70.03
-13.1 | [5 empty] | 18.78 | 70.43
-13 | [5 empty] | 19.29 | 70.73
-12.9 |  |  |  |  | 5.4 | 19.79 | 71.2
-12.8 |  |  |  |  | 5.49 | 20.22 | 71.56
-12.7 |  |  |  |  | 5.59 | 21.07 | 71.74
-12.6 |  |  |  |  | 5.55 | 21.88 | 72.21
-12.5 |  |  |  |  | 5.81 | 23.63 | 72.54
-12.4 |  |  |  |  | 5.88 | 25.32 | 72.92
-12.3 |  |  |  |  | 5.82 | 28.32 | 73.25
-12.2 |  |  |  |  | 5.97 | 32.31 | 73.54
-12.1 |  |  |  |  | 6.08 | 37.63 | 73.84
-12 |  |  |  |  | 6.1 | 41.78 | 74.13
-11.9 |  |  |  |  | 6.26 | 48.44 | 74.53
-11.8 |  |  |  |  | 6.38 | 53.42 | 74.82
-11.7 |  |  |  | 4.19 | 6.4 | 57.31 | 75.32
-11.6 |  |  |  | 4.19 | 6.57 | 61.66 | 75.48
-11.5 |  |  |  | 4.21 | 6.69 | 64.22 | 75.8
-11.4 |  |  |  | 4.32 | 6.78 | 65.9 | 76.24
-11.3 |  |  |  | 4.41 | 6.87 | 67.55 | 76.44
-11.2 |  |  |  | 4.52 | 6.96 | 68.48 | 76.9
-11.1 |  |  |  | 4.6 | 7.2 | 69.3 | 77.19
-11 |  |  |  | 4.59 | 7.28 | 70.05 | 77.34
-10.9 |  |  |  | 4.61 | 7.33 | 70.77 | 77.81
-10.8 |  |  |  | 4.71 | 7.47 | 71.24 | 77.94
-10.7 |  |  |  | 4.84 | 7.71 | 71.89 | 78.3
-10.6 |  |  |  | 5.01 | 7.78 | 72.45 | 78.8
-10.5 |  |  | 3.3 | 5 | 7.88 | 72.86 | 79.06
-10.4 |  |  | 3.3 | 5.12 | 7.98 | 73.4 | 79.36
-10.3 |  |  | 3.3 | 5.21 | 8.12 | 74.11 | 79.63
-10.2 |  |  | 3.46 | 5.31 | 8.29 | 74.26 | 79.85
-10.1 |  |  | 3.58 | 5.41 | 8.5 | 74.88 | 80.32
-10 |  |  | 3.52 | 5.51 | 8.58 | 75.26 | 80.63
-9.9 |  |  | 3.5 | 5.61 | 8.75 | 75.72 | 80.94
-9.8 |  |  | 3.58 | 5.71 | 8.85 | 76.21 | 81.11
-9.7 |  |  | 3.63 | 5.81 | 9.09 | 76.71 | 81.52
-9.6 |  |  | 3.82 | 5.92 | 9.15 | 76.92 | 81.88
-9.5 |  |  | 3.8 | 6.01 | 9.37 | 77.45 | 82.01
-9.4 |  |  | 3.88 | 6.14 | 9.67 | 77.93 | 82.34
-9.3 |  | 1.97 | 3.98 | 6.3 | 9.78 | 78.12 | 82.69
-9.2 |  | 1.92 | 4 | 6.32 | 9.96 | 78.62 | 83.03
-9.1 |  | 2.02 | 4.08 | 6.53 | 10.15 | 79.1 | 83.29
-9 |  | 1.94 | 4.26 | 6.61 | 10.32 | 79.6 | 83.55
-8.9 |  | 2.09 | 4.22 | 6.73 | 10.56 | 79.81 | 83.89
-8.8 |  | 2.09 | 4.28 | 6.92 | 10.69 | 80.18 | 84.15
-8.7 |  | 2.09 | 4.38 | 7.03 | 10.94 | 80.7 | 84.42
-8.6 |  | 2.19 | 4.48 | 7.22 | 11.2 | 81.17 | 84.72
-8.5 |  | 2.14 | 4.5 | 7.31 | 11.36 | 81.34 | 84.91
-8.4 |  | 2.29 | 4.58 | 7.43 | 11.52 | 81.75 | 85.28
-8.3 |  | 2.29 | 4.68 | 7.61 | 11.84 | 82.3 | 85.39
-8.2 |  | 2.34 | 4.73 | 7.74 | 12.14 | 82.72 | 85.93
-8.1 |  | 2.34 | 4.92 | 8.03 | 12.47 | 83.03 | 85.91
-8 |  | 2.39 | 5.06 | 8.12 | 12.63 | 83.25 | 86.49
-7.9 |  | 2.44 | 5.09 | 8.33 | 12.94 | 83.74 | 86.72
-7.8 |  | 2.49 | 5.05 | 8.53 | 13.24 | 83.96 | 86.82
-7.7 |  | 2.55 | 5.38 | 8.71 | 13.54 | 84.28 | 87.39
-7.6 |  | 2.49 | 5.37 | 8.82 | 13.84 | 84.72 | 87.62
-7.5 |  | 2.6 | 5.47 | 9.03 | 14.22 | 85.05 | 87.77
-7.4 |  | 2.68 | 5.57 | 9.25 | 14.51 | 85.44 | 87.89
-7.3 |  | 2.74 | 5.68 | 9.52 | 15.03 | 85.78 | 88.51
-7.2 |  | 2.79 | 5.84 | 9.61 | 15.32 | 86.2 | 88.43
-7.1 |  | 2.74 | 5.97 | 9.85 | 15.91 | 86.46 | 88.93
-7 |  | 2.89 | 6.06 | 10.13 | 16.37 | 86.71 | 88.98
-6.9 |  | 3.03 | 6.16 | 10.34 | 17.06 | 87.19 | 89.51
-6.8 |  | 3.03 | 6.28 | 10.64 | 17.81 | 87.41 | 89.61
-6.7 |  | 3.09 | 6.59 | 10.82 | 18.86 | 87.83 | 89.88
-6.6 |  | 3.23 | 6.48 | 11.06 | 20.33 | 88.13 | 90.08
-6.5 |  | 3.29 | 6.78 | 11.47 | 22.62 | 88.39 | 90.35
-6.4 |  | 3.38 | 6.95 | 11.72 | 25.96 | 88.6 | 90.61
-6.3 |  | 3.38 | 7.2 | 11.98 | 30.3 | 89.03 | 90.84
-6.2 |  | 3.36 | 7.14 | 12.46 | 36.92 | 89.21 | 91.1
-6.1 |  | 3.55 | 7.35 | 12.77 | 43.42 | 89.69 | 91.41
-6 |  | 3.6 | 7.74 | 13.28 | 51.71 | 89.84 | 91.67
-5.9 |  | 3.73 | 7.92 | 13.67 | 58.74 | 90.17 | 91.81
-5.8 |  | 3.68 | 7.94 | 14.21 | 65.91 | 90.41 | 92.11
-5.7 |  | 3.88 | 8.21 | 14.93 | 71.99 | 90.8 | 92.39
-5.6 |  | 3.91 | 8.44 | 15.66 | 76.49 | 91.06 | 92.33
-5.5 |  | 4.11 | 8.65 | 16.73 | 79.98 | 91.36 | 92.88
-5.4 |  | 4.14 | 9.03 | 18.55 | 82.43 | 91.59 | 92.96
-5.3 |  | 4.28 | 9.18 | 20.76 | 83.89 | 91.74 | 93.12
-5.2 |  | 4.49 | 9.57 | 24.41 | 85.22 | 92.19 | 93.48
-5.1 | 1.3 | 4.4 | 9.8 | 29.56 | 86.2 | 92.34 | 93.66
-5 | 1.2 | 4.58 | 9.98 | 36.18 | 86.96 | 92.73 | 93.92
-4.9 | 1.28 | 4.65 | 10.44 | 43.87 | 87.61 | 92.96 | 94.02
-4.8 | 1.37 | 4.8 | 10.81 | 52.3 | 88.27 | 93.13 | 94.36
-4.7 | 1.34 | 5.12 | 11.25 | 60.67 | 88.83 | 93.47 | 94.6
-4.6 | 1.4 | 5.07 | 11.81 | 68.45 | 89.4 | 93.71 | 94.8
-4.5 | 1.44 | 5.3 | 12.41 | 74.81 | 89.8 | 94.03 | 94.87
-4.4 | 1.46 | 5.49 | 13.1 | 79.85 | 90.36 | 94.19 | 95.03
-4.3 | 1.43 | 5.45 | 14.25 | 83.31 | 90.81 | 94.45 | 95.42
-4.2 | 1.56 | 5.86 | 15.84 | 85.84 | 91.39 | 94.63 | 95.53
-4.1 | 1.61 | 6.14 | 18.19 | 87.66 | 91.65 | 95.02 | 95.83
-4 | 1.7 | 6.1 | 22.1 | 88.82 | 92.07 | 95.12 | 95.99
-3.9 | 1.82 | 6.5 | 27.44 | 89.8 | 92.47 | 95.3 | 96.17
-3.8 | 1.67 | 6.69 | 34.31 | 90.58 | 92.87 | 95.6 | 96.31
-3.7 | 1.92 | 7 | 42.53 | 91.27 | 93.22 | 95.8 | 96.52
-3.6 | 1.9 | 7.2 | 51.66 | 91.86 | 93.63 | 96 | 96.75
-3.5 | 1.94 | 7.36 | 60.84 | 92.36 | 94.04 | 96.21 | 96.77
-3.4 | 1.99 | 7.95 | 69.3 | 92.85 | 94.3 | 96.46 | 97.03
-3.3 | 2.13 | 8.16 | 76.61 | 93.26 | 94.63 | 96.54 | 97.29
-3.2 | 2.18 | 8.66 | 82.22 | 93.85 | 95 | 96.86 | 97.34
-3.1 | 2.13 | 9.51 | 85.6 | 94.14 | 95.31 | 96.98 | 97.52
-3 | 2.38 | 10.15 | 89.35 | 94.66 | 95.56 | 97.17 | 97.68
-2.9 | 2.56 | 11.64 | 91.01 | 95.14 | 95.88 | 97.32 | 97.79
-2.8 | 2.74 | 13.93 | 92.26 | 95.43 | 96.24 | 97.38 | 97.97
-2.7 | 2.71 | 18.27 | 93.19 | 95.63 | 96.39 | 97.53 | 98.1
-2.6 | 2.6 | 26.97 | 94 | 95.94 | 96.71 | 97.87 | 98.35
-2.5 | 2.81 | 38.89 | 94.71 | 96.34 | 97.01 | 98.04 | 98.45
-2.4 | 3.49 | 53.73 | 95.19 | 96.63 | 97.25 | 98.17 | 98.51
-2.3 | 3.34 | 67.64 | 95.77 | 96.93 | 97.51 | 98.31 | 98.55
-2.2 | 3.73 | 79.06 | 96.06 | 97.13 | 97.72 | 98.35 | 98.68
-2.1 | 3.95 | 86.54 | 96.35 | 97.43 | 97.73 | 98.65 | 98.79
-2 | 4.32 | 90.39 | 96.8 | 97.7 | 98.02 | 98.71 | 98.97
-1.9 | 4.84 | 93.62 | 97.05 | 97.91 | 98.34 | 98.8 | 99.07
-1.8 | 5.62 | 94.36 | 97.25 | 98.03 | 98.46 | 98.9 | 99.09
-1.7 | 6.99 | 95.1 | 97.7 | 98.43 | 98.62 | 99 | 99.16
-1.6 | 9.57 | 96.03 | 97.9 | 98.5 | 98.73 | 99.06 | 99.21
-1.5 | 14.64 | 96.41 | 98.05 | 98.62 | 99.04 | 99.16 | 99.32
-1.4 | 23.5 | 97 | 98.4 | 98.8 | 99.16 | 99.3 | 99.46
-1.3 | 37.7 | 97.49 | 98.65 | 98.9 | 99.26 | 99.4 | 99.5
-1.2 | 54.42 | 97.88 | 98.9 | 99.1 | 99.36 | 99.4 | 99.62
-1.1 | 69.2 | 98.34 | 99.15 | 99.4 | 99.43 | 99.36 | 99.69
-1 | 81.3 | 98.5 | 99.15 | 99.3 | 99.68 | 99.65 | 99.77
-0.9 | 89.4 | 98.58 | 99.4 | 99.4 | 99.7 | 99.56 | 99.74
-0.8 | 94.3 | 99.13 | 99.45 | 99.6 | 99.65 | 99.7 | 99.82
-0.7 | 96.7 | 98.79 | 99.69 | 99.7 | 99.72 | 99.61 | 99.84
-0.6 | 97.8 | 99.24 | 99.51 | 99.7 | 99.8 | 99.8 | 99.86
-0.5 | 99.1 | 99.4 | 99.8 | 99.8 | 99.85 | 100 | 99.9
-0.4 | 99.7 | 99.69 | 99.85 | 100 | 99.9 | 99.91 | 99.96
-0.3 | 99.7 | 99.8 | 99.8 | 99.9 | 99.96 | 99.85 | 100
-0.2 | 99.9 | 100 | 100 | 100 | 100 | 99.85 | 100
-0.1 | 100 | 99.3 | 99.95 | 99.9 | 100 | 99.9 | 100
0 | 100 | 99.5 | 99.95 | 99.9 | 100 | 99.95 | 100
0.1 | 100 | 99.3 | 99.95 | 99.9 | 100 | 99.9 | 100
0.2 | 99.9 | 100 | 100 | 100 | 100 | 99.85 | 100
0.3 | 99.7 | 99.8 | 99.8 | 99.9 | 99.94 | 99.85 | 100
0.4 | 99.7 | 99.7 | 99.85 | 100 | 99.9 | 99.91 | 99.96
0.5 | 99.1 | 99.4 | 99.8 | 99.8 | 99.85 | 100 | 99.9
0.6 | 97.8 | 99.2 | 99.51 | 99.7 | 99.8 | 99.8 | 99.85
0.7 | 96.7 | 98.78 | 99.69 | 99.7 | 99.72 | 99.61 | 99.84
0.8 | 94.3 | 99.1 | 99.45 | 99.6 | 99.65 | 99.7 | 99.82
0.9 | 89.4 | 98.51 | 99.4 | 99.4 | 99.7 | 99.56 | 99.74
1 | 81.3 | 98.5 | 99.15 | 99.3 | 99.68 | 99.65 | 99.77
1.1 | 69.2 | 98.34 | 99.15 | 99.4 | 99.43 | 99.36 | 99.69
1.2 | 54.42 | 97.88 | 98.9 | 99.1 | 99.36 | 99.4 | 99.62
1.3 | 37.7 | 97.49 | 98.65 | 98.9 | 99.26 | 99.4 | 99.5
1.4 | 23.5 | 97 | 98.4 | 98.8 | 99.16 | 99.3 | 99.46
1.5 | 14.64 | 96.41 | 98.05 | 98.62 | 99.04 | 99.16 | 99.32
1.6 | 9.57 | 96.03 | 97.9 | 98.5 | 98.73 | 99.06 | 99.21
1.7 | 6.99 | 95.1 | 97.7 | 98.43 | 98.62 | 99 | 99.16
1.8 | 5.62 | 94.36 | 97.25 | 98.03 | 98.46 | 98.9 | 99.09
1.9 | 4.84 | 93.62 | 97.05 | 97.91 | 98.34 | 98.8 | 99.07
2 | 4.32 | 90.39 | 96.8 | 97.7 | 98.02 | 98.71 | 98.97
2.1 | 3.95 | 86.54 | 96.35 | 97.43 | 97.73 | 98.65 | 98.79
2.2 | 3.73 | 79.07 | 96.05 | 97.13 | 97.72 | 98.36 | 98.68
2.3 | 3.34 | 67.83 | 95.75 | 96.93 | 97.51 | 98.3 | 98.55
2.4 | 3.49 | 53.73 | 95.21 | 96.63 | 97.25 | 98.17 | 98.51
2.5 | 2.81 | 38.89 | 94.67 | 96.34 | 97.01 | 98.06 | 98.45
2.6 | 2.6 | 26.97 | 94.02 | 95.94 | 96.71 | 97.86 | 98.35
2.7 | 2.71 | 18.27 | 93.11 | 95.63 | 96.39 | 97.55 | 98.1
2.8 | 2.74 | 13.93 | 92.26 | 95.43 | 96.24 | 97.35 | 97.97
2.9 | 2.56 | 11.64 | 91.01 | 95.14 | 95.88 | 97.31 | 97.79
3 | 2.38 | 10.15 | 89.35 | 94.66 | 95.56 | 97.17 | 97.68
3.1 | 2.13 | 9.51 | 85.63 | 94.14 | 95.31 | 96.98 | 97.52
3.2 | 2.18 | 8.66 | 82.67 | 93.85 | 94.97 | 96.9 | 97.34
3.3 | 2.13 | 8.16 | 75.56 | 93.26 | 94.65 | 96.45 | 97.28
3.4 | 1.99 | 7.95 | 69.26 | 92.85 | 94.31 | 96.46 | 97.01
3.5 | 1.94 | 7.36 | 60.84 | 92.36 | 94.04 | 96.21 | 96.78
3.6 | 1.9 | 7.2 | 51.66 | 91.86 | 93.63 | 96 | 96.74
3.7 | 1.92 | 7 | 42.53 | 91.27 | 93.22 | 95.8 | 96.52
3.8 | 1.67 | 6.69 | 34.31 | 90.58 | 92.87 | 95.6 | 96.31
3.9 | 1.82 | 6.5 | 27.44 | 89.8 | 92.47 | 95.3 | 96.17
4 | 1.7 | 6.09 | 22.1 | 88.82 | 92.07 | 95.12 | 95.99
4.1 | 1.61 | 6.14 | 18.19 | 87.66 | 91.65 | 95.02 | 95.8
4.2 | 1.56 | 5.86 | 15.84 | 85.84 | 91.39 | 94.63 | 95.53
4.3 | 1.43 | 5.45 | 14.25 | 83.31 | 90.81 | 94.45 | 95.42
4.4 | 1.46 | 5.49 | 13.1 | 79.85 | 90.36 | 94.19 | 95.02
4.5 | 1.44 | 5.3 | 12.41 | 74.81 | 89.8 | 94.03 | 94.86
4.6 | 1.4 | 5.07 | 11.81 | 68.45 | 89.4 | 93.71 | 94.79
4.7 | 1.34 | 5.12 | 11.25 | 60.67 | 88.8 | 93.47 | 94.58
4.8 | 1.37 | 4.8 | 10.81 | 52.3 | 88.28 | 93.13 | 94.32
4.9 | 1.28 | 4.65 | 10.44 | 43.87 | 87.61 | 92.96 | 94.02
5 | 1.2 | 4.58 | 9.98 | 36.18 | 86.96 | 92.73 | 93.89
5.1 |  | 4.4 | 9.77 | 29.56 | 86.2 | 92.34 | 93.66
5.2 |  | 4.53 | 9.57 | 24.41 | 85.22 | 92.19 | 93.46
5.3 |  | 4.28 | 9.18 | 20.77 | 83.89 | 91.74 | 93.09
5.4 |  | 4.14 | 9.03 | 18.38 | 82.43 | 91.59 | 92.96
5.5 |  | 4.11 | 8.65 | 16.71 | 79.98 | 91.36 | 92.88
5.6 |  | 3.9 | 8.44 | 15.67 | 76.49 | 91.06 | 92.33
5.7 |  | 3.88 | 8.21 | 14.86 | 71.99 | 90.8 | 92.39
5.8 |  | 3.68 | 7.94 | 14.21 | 65.9 | 90.41 | 92.11
5.9 |  | 3.71 | 7.92 | 13.63 | 58.52 | 90.17 | 91.81
6 |  | 3.6 | 7.74 | 13.27 | 51.27 | 89.84 | 91.67
6.1 |  | 3.55 | 7.35 | 12.77 | 43.66 | 89.69 | 91.41
6.2 |  | 3.37 | 7.14 | 12.46 | 36.15 | 89.21 | 91.1
6.3 |  | 3.38 | 7.2 | 11.98 | 30.27 | 89.03 | 90.84
6.4 |  | 3.38 | 6.95 | 11.72 | 25.95 | 88.6 | 90.61
6.5 |  | 3.29 | 6.74 | 11.47 | 22.62 | 88.39 | 90.35
6.6 |  | 3.23 | 6.54 | 11.06 | 20.36 | 88.13 | 90.08
6.7 |  | 3.09 | 6.59 | 10.82 | 18.7 | 87.83 | 89.88
6.8 |  | 3.03 | 6.28 | 10.64 | 17.82 | 87.41 | 89.61
6.9 |  | 3.03 | 6.16 | 10.34 | 16.97 | 87.19 | 89.51
7 |  | 2.89 | 6.06 | 10.13 | 16.37 | 86.71 | 88.98
7.1 |  | 2.74 | 5.96 | 9.85 | 15.91 | 86.46 | 88.93
7.2 |  | 2.79 | 5.83 | 9.61 | 15.32 | 86.2 | 88.43
7.3 |  | 2.74 | 5.67 | 9.52 | 15.03 | 85.78 | 88.51
7.4 |  | 2.68 | 5.56 | 9.25 | 14.51 | 85.44 | 87.89
7.5 |  | 2.6 | 5.46 | 9.03 | 14.22 | 85.05 | 87.77
7.6 |  | 2.49 | 5.37 | 8.82 | 13.84 | 84.72 | 87.62
7.7 |  | 2.55 | 5.33 | 8.71 | 13.54 | 84.28 | 87.39
7.8 |  | 2.49 | 5.11 | 8.53 | 13.24 | 83.96 | 86.82
7.9 |  | 2.44 | 5.09 | 8.33 | 12.94 | 83.71 | 86.72
8 |  | 2.39 | 5.06 | 8.12 | 12.63 | 83.3 | 86.49
8.1 |  | 2.34 | 4.95 | 8.03 | 12.47 | 83.03 | 85.91
8.2 |  | 2.34 | 4.64 | 7.74 | 12.14 | 82.69 | 85.94
8.3 |  | 2.29 | 4.67 | 7.61 | 11.84 | 82.29 | 85.39
8.4 |  | 2.29 | 4.58 | 7.43 | 11.52 | 81.81 | 85.31
8.5 |  | 2.14 | 4.5 | 7.31 | 11.36 | 81.35 | 84.82
8.6 |  | 2.19 | 4.48 | 7.22 | 11.2 | 81.17 | 84.72
8.7 |  | 2.09 | 4.38 | 7.03 | 10.94 | 80.65 | 84.42
8.8 |  | 2.09 | 4.28 | 6.92 | 10.69 | 80.24 | 84.12
8.9 |  | 2.09 | 4.22 | 6.73 | 10.56 | 79.81 | 83.92
9 |  | 1.94 | 4.26 | 6.61 | 10.32 | 79.59 | 83.47
9.1 |  | 2.05 | 4.08 | 6.53 | 10.15 | 79.05 | 83.29
9.2 |  | 1.94 | 4 | 6.32 | 9.96 | 78.64 | 83.03
9.3 |  |  | 3.98 | 6.3 | 9.78 | 78.18 | 82.69
9.4 |  |  | 3.88 | 6.14 | 9.67 | 77.93 | 82.34
9.5 |  |  | 3.8 | 6.01 | 9.37 | 77.45 | 82.01
9.6 |  |  | 3.82 | 5.92 | 9.15 | 76.92 | 81.88
9.7 |  |  | 3.63 | 5.81 | 9.09 | 76.71 | 81.52
9.8 |  |  | 3.58 | 5.71 | 8.85 | 76.21 | 81.11
9.9 |  |  | 3.5 | 5.61 | 8.75 | 75.72 | 80.94
10 |  |  | 3.52 | 5.51 | 8.58 | 75.26 | 80.63
10.1 |  |  | 3.58 | 5.41 | 8.5 | 74.88 | 80.32
10.2 |  |  | 3.46 | 5.31 | 8.29 | 74.26 | 79.85
10.3 |  |  | 3.3 | 5.21 | 8.12 | 74.11 | 79.63
10.4 |  |  | 3.3 | 5.12 | 7.98 | 73.4 | 79.36
10.5 |  |  |  | 5 | 7.88 | 72.86 | 79.06
10.6 |  |  |  | 5.01 | 7.78 | 72.45 | 78.8
10.7 |  |  |  | 4.84 | 7.71 | 71.89 | 78.3
10.8 |  |  |  | 4.71 | 7.47 | 71.24 | 77.94
10.9 |  |  |  | 4.61 | 7.33 | 70.77 | 77.81
11 |  |  |  | 4.59 | 7.28 | 70.05 | 77.35
11.1 |  |  |  | 4.6 | 7.2 | 69.3 | 77.18
11.2 |  |  |  | 4.52 | 6.96 | 68.48 | 76.86
11.3 |  |  |  | 4.41 | 6.87 | 67.55 | 76.45
11.4 |  |  |  | 4.32 | 6.78 | 65.9 | 76.23
11.5 |  |  |  | 4.21 | 6.69 | 64.22 | 75.76
11.6 |  |  |  | 4.19 | 6.57 | 61.66 | 75.48
11.7 |  |  |  |  | 6.4 | 57.31 | 75.32
11.8 |  |  |  |  | 6.38 | 53.42 | 74.82
11.9 |  |  |  |  | 6.26 | 48.44 | 74.53
12 |  |  |  |  | 6.1 | 41.78 | 74.13
12.1 |  |  |  |  | 6.08 | 37.63 | 73.84
12.2 |  |  |  |  | 5.97 | 32.31 | 73.54
12.3 |  |  |  |  | 5.82 | 28.32 | 73.25
12.4 |  |  |  |  | 5.88 | 25.32 | 72.92
12.5 |  |  |  |  | 5.81 | 23.63 | 72.54
12.6 |  |  |  |  | 5.55 | 21.88 | 72.21
12.7 |  |  |  |  | 5.59 | 21.07 | 71.74
12.8 |  |  |  |  | 5.49 | 20.22 | 71.56
12.9 | [5 empty] | 19.79 | 71.2
13 | [5 empty] | 19.29 | 70.73
13.1 | [5 empty] | 18.78 | 70.43
13.2 | [5 empty] | 18.32 | 70.03
13.3 | [5 empty] | 18.03 | 69.75
13.4 | [5 empty] | 17.6 | 69.43
13.5 | [5 empty] | 17.37 | 69.01
13.6 | [5 empty] | 17 | 68.64
13.7 | [5 empty] | 16.75 | 68.33
13.8 | [5 empty] | 16.38 | 67.92
13.9 | [5 empty] | 15.96 | 67.53
14 | [5 empty] | 15.83 | 67.27
14.1 | [5 empty] | 15.55 | 67.02
14.2 | [5 empty] | 15.31 | 66.51
14.3 | [5 empty] | 14.85 | 66.12
14.4 | [5 empty] | 14.75 | 65.72
14.5 | [5 empty] | 14.47 | 65.32
14.6 | [5 empty] | 14.31 | 64.92
14.7 | [5 empty] | 13.87 | 64.52
14.8 | [5 empty] | 13.86 | 64.13
14.9 | [5 empty] | 13.68 | 63.84
15 | [5 empty] | 13.43 | 63.38
15.1 | [5 empty] | 13.28 | 62.82
15.2 | [5 empty] | 13.03 | 62.52
15.3 | [5 empty] | 12.87 | 62.01
15.4 | [5 empty] | 12.65 | 61.6
15.5 | [5 empty] | 12.39 | 61.14
15.6 | [5 empty] | 12.26 | 60.8
15.7 | [5 empty] | 12.05 | 60.2
15.8 | [5 empty] | 11.85 | 59.8
15.9 | [5 empty] | 11.66 | 59.16
16 | [5 empty] | 11.45 | 58.48
16.1 | [5 empty] | 11.27 | 57.83
16.2 | [5 empty] | 11.18 | 56.78
16.3 | [5 empty] | 11.06 | 55.51
16.4 | [5 empty] | 10.85 | 53.66
16.5 | [5 empty] | 10.66 | 51.03
16.6 | [5 empty] | 10.58 | 47.8
16.7 | [5 empty] | 10.36 | 43.84
16.8 | [5 empty] | 10.25 | 39.46
16.9 | [5 empty] | 10.06 | 35.34
17 | [5 empty] | 10.02 | 31.11
17.1 | [5 empty] | 9.66 | 27.67
17.2 | [5 empty] | 9.69 | 24.9
17.3 | [5 empty] | 9.46 | 22.87
17.4 | [5 empty] | 9.36 | 21.56
17.5 | [5 empty] | 9.29 | 20.53
17.6 | [5 empty] | 9.14 | 19.88
17.7 | [5 empty] | 8.97 | 19.29
17.8 | [5 empty] | 8.88 | 18.81
17.9 | [5 empty] | 8.77 | 18.44
18 | [5 empty] | 8.69 | 18.12
18.1 | [5 empty] | 8.54 | 17.73
18.2 | [5 empty] | 8.38 | 17.44
18.3 | [5 empty] | 8.29 | 17.14
18.4 | [5 empty] | 8.16 | 16.84
18.5 | [5 empty] | 8.09 | 16.56
18.6 | [5 empty] | 7.96 | 16.35
18.7 | [5 empty] | 7.88 | 16.03
18.8 | [5 empty] | 7.8 | 15.76
18.9 | [6 empty] | 15.55
19 | [6 empty] | 15.28
19.1 | [6 empty] | 15.08
19.2 | [6 empty] | 14.87
19.3 | [6 empty] | 14.57
19.4 | [6 empty] | 14.38
19.5 | [6 empty] | 14.18
19.6 | [6 empty] | 13.98
19.7 | [6 empty] | 13.79
19.8 | [6 empty] | 13.68
19.9 | [6 empty] | 13.37
20 | [6 empty] | 13.19
20.1 | [6 empty] | 13.08
20.2 | [6 empty] | 12.77
20.3 | [6 empty] | 12.59
20.4 | [6 empty] | 12.49
20.5 | [6 empty] | 12.38
20.6 | [6 empty] | 12.07
20.7 | [6 empty] | 11.89
20.8 | [6 empty] | 11.79
20.9 | [6 empty] | 11.68
21 | [6 empty] | 11.49
21.1 | [6 empty] | 11.38
21.2 | [6 empty] | 11.08
21.3 | [6 empty] | 10.99
21.4 | [6 empty] | 10.89
21.5 | [6 empty] | 10.78
21.6 | [6 empty] | 10.59
21.7 | [6 empty] | 10.48
21.8 | [6 empty] | 10.28
21.9 | [6 empty] | 10.19
22 | [6 empty] | 10.08
22.1 | [6 empty] | 9.88
22.2 | [6 empty] | 9.79
22.3 | [6 empty] | 9.69
22.4 | [6 empty] | 9.59
22.5 | [6 empty] | 9.48
22.6 | [6 empty] | 9.28
22.7
22.8
22.9
23
23.1
23.2
23.3
23.4
23.5
23.6
23.7
23.8
23.9
24
24.1
24.2
24.3
24.4
24.5
24.6
24.7
24.8
24.9
25
25.1
25.2

## Diagonal Profiles
Diagonal Profiles
Open field
SSD = 90 cm
Field Size = 28x28 cm2
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.3 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-32.2 |  |  |  |  | 3.61
-32.1 |  |  |  |  | 3.51
-32 |  |  |  |  | 3.51
-31.9 |  |  |  |  | 3.5
-31.8 |  |  |  |  | 3.59
-31.7 |  |  |  |  | 3.62
-31.6 |  |  |  | 2.04 | 3.61
-31.5 |  |  |  | 2.08 | 3.62
-31.4 |  |  |  | 2.19 | 3.59
-31.3 |  |  |  | 2.2 | 3.79
-31.2 |  |  |  | 2.19 | 3.81
-31.1 |  |  |  | 2.29 | 3.89
-31 |  |  |  | 2.3 | 3.92
-30.9 |  |  |  | 2.3 | 3.9
-30.8 |  |  |  | 2.3 | 3.98
-30.7 |  |  |  | 2.29 | 4.1
-30.6 |  |  |  | 2.39 | 4.12
-30.5 |  |  |  | 2.4 | 4.11
-30.4 |  |  |  | 2.49 | 4.18
-30.3 |  |  |  | 2.51 | 4.3
-30.2 |  |  |  | 2.49 | 4.31
-30.1 |  |  |  | 2.58 | 4.4
-30 |  |  |  | 2.6 | 4.42
-29.9 |  |  |  | 2.6 | 4.41
-29.8 |  |  |  | 2.59 | 4.48
-29.7 |  |  |  | 2.69 | 4.58
-29.6 |  |  | 1.27 | 2.69 | 4.7
-29.5 |  |  | 1.3 | 2.8 | 4.71
-29.4 |  |  | 1.3 | 2.71 | 4.79
-29.3 |  |  | 1.3 | 2.77 | 4.81
-29.2 |  |  | 1.3 | 2.8 | 4.88
-29.1 |  |  | 1.3 | 2.88 | 5
-29 |  |  | 1.32 | 2.9 | 5.01
-28.9 |  |  | 1.41 | 2.98 | 5.08
-28.8 |  |  | 1.4 | 3.01 | 5.18
-28.7 |  |  | 1.4 | 2.99 | 5.29
-28.6 |  | 0.9 | 1.4 | 3.09 | 5.32
-28.5 |  | 0.9 | 1.4 | 3.1 | 5.38
-28.4 |  | 0.9 | 1.41 | 3.17 | 5.5
-28.3 |  | 0.9 | 1.51 | 3.21 | 5.51
-28.2 |  | 0.9 | 1.5 | 3.19 | 5.58
-28.1 |  | 0.9 | 1.5 | 3.28 | 5.7
-28 |  | 0.9 | 1.5 | 3.3 | 5.71
-27.9 | 0.7 | 0.9 | 1.5 | 3.37 | 5.79
-27.8 | 0.7 | 1 | 1.51 | 3.48 | 5.89
-27.7 | 0.7 | 1 | 1.6 | 3.51 | 5.98
-27.6 | 0.7 | 1 | 1.6 | 3.49 | 6.08
-27.5 | 0.7 | 1 | 1.6 | 3.58 | 6.18
-27.4 | 0.7 | 1 | 1.6 | 3.6 | 6.3
-27.3 | 0.7 | 1 | 1.61 | 3.67 | 6.32
-27.2 | 0.7 | 1 | 1.71 | 3.7 | 6.38
-27.1 | 0.7 | 1 | 1.7 | 3.76 | 6.49
-27 | 0.7 | 1 | 1.71 | 3.88 | 6.59
-26.9 | 0.7 | 1.1 | 1.81 | 3.91 | 6.68
-26.8 | 0.71 | 1.1 | 1.8 | 3.89 | 6.87
-26.7 | 0.8 | 1.1 | 1.8 | 3.96 | 6.99
-26.6 | 0.8 | 1.1 | 1.8 | 4.08 | 7.08
-26.5 | 0.8 | 1.1 | 1.9 | 4.1 | 7.18
-26.4 | 0.8 | 1.1 | 1.9 | 4.17 | 7.28
-26.3 | 0.8 | 1.1 | 1.9 | 4.2 | 7.39
-26.2 | 0.8 | 1.2 | 1.91 | 4.26 | 7.48
-26.1 | 0.8 | 1.2 | 2.01 | 4.36 | 7.65
-26 | 0.8 | 1.2 | 2 | 4.47 | 7.8
-25.9 | 0.8 | 1.2 | 2 | 4.5 | 7.89
-25.8 | 0.8 | 1.2 | 2.01 | 4.56 | 7.99
-25.7 | 0.81 | 1.3 | 2.1 | 4.67 | 8.14
-25.6 | 0.91 | 1.3 | 2.1 | 4.78 | 8.29
-25.5 | 0.9 | 1.3 | 2.1 | 4.8 | 8.37
-25.4 | 0.9 | 1.3 | 2.2 | 4.85 | 8.55
-25.3 | 0.9 | 1.3 | 2.2 | 4.96 | 8.68
-25.2 | 0.9 | 1.3 | 2.3 | 5.07 | 8.85
-25.1 | 0.9 | 1.29 | 2.3 | 5.1 | 8.99
-25 | 0.9 | 1.39 | 2.3 | 5.15 | 9.16
-24.9 | 0.9 | 1.4 | 2.41 | 5.26 | 9.28
-24.8 | 0.9 | 1.4 | 2.4 | 5.36 | 9.51
-24.7 | 0.9 | 1.4 | 2.4 | 5.46 | 9.76
-24.6 | 1 | 1.4 | 2.5 | 5.57 | 9.92
-24.5 | 1 | 1.5 | 2.5 | 5.66 | 10.18
-24.4 | 1 | 1.5 | 2.5 | 5.75 | 10.52
-24.3 | 1 | 1.5 | 2.6 | 5.85 | 11.04
-24.2 | 1 | 1.5 | 2.7 | 5.95 | 11.69
-24.1 | 1 | 1.49 | 2.7 | 6.06 | 12.59
-24 | 1.01 | 1.59 | 2.8 | 6.15 | 13.94
-23.9 | 1.1 | 1.61 | 2.8 | 6.32 | 15.69
-23.8 | 1.1 | 1.59 | 2.9 | 6.47 | 17.82
-23.7 | 1.1 | 1.69 | 2.9 | 6.56 | 20.15
-23.6 | 1.1 | 1.7 | 3 | 6.67 | 22.93
-23.5 | 1.1 | 1.7 | 3 | 6.74 | 25.97
-23.4 | 1.1 | 1.7 | 3.1 | 6.9 | 28.86
-23.3 | 1.2 | 1.8 | 3.1 | 7.05 | 31.57
-23.2 | 1.2 | 1.8 | 3.19 | 7.19 | 33.96
-23.1 | 1.2 | 1.8 | 3.3 | 7.4 | 35.95
-23 | 1.2 | 1.79 | 3.3 | 7.61 | 37.44
-22.9 | 1.2 | 1.89 | 3.4 | 7.75 | 38.57
-22.8 | 1.21 | 1.9 | 3.4 | 7.9 | 39.54
-22.7 | 1.3 | 1.99 | 3.49 | 8.11 | 40.28
-22.6 | 1.3 | 2.01 | 3.59 | 8.34 | 40.86
-22.5 | 1.3 | 1.99 | 3.7 | 8.61 | 41.51
-22.4 | 1.3 | 2.09 | 3.7 | 9 | 42.03
-22.3 | 1.3 | 2.11 | 3.79 | 9.55 | 42.41
-22.2 | 1.4 | 2.1 | 3.89 | 10.27 | 42.82
-22.1 | 1.4 | 2.2 | 3.99 | 11.36 | 43.22
-22 | 1.4 | 2.2 | 4.1 | 12.89 | 43.64
-21.9 | 1.4 | 2.29 | 4.2 | 15.01 | 43.99
-21.8 | 1.41 | 2.3 | 4.3 | 17.69 | 44.33
-21.7 | 1.5 | 2.39 | 4.4 | 20.89 | 44.75
-21.6 | 1.5 | 2.4 | 4.5 | 24.11 | 45.14
-21.5 | 1.5 | 2.48 | 4.59 | 27.49 | 45.47
-21.4 | 1.5 | 2.59 | 4.68 | 31 | 45.76
-21.3 | 1.6 | 2.6 | 4.88 | 34 | 46.11
-21.2 | 1.6 | 2.69 | 4.99 | 36.55 | 46.49
-21.1 | 1.6 | 2.7 | 5.08 | 38.52 | 46.78
-21 | 1.7 | 2.77 | 5.28 | 40.1 | 47.08
-20.9 | 1.7 | 2.89 | 5.39 | 41.12 | 47.42
-20.8 | 1.7 | 2.9 | 5.58 | 41.99 | 47.85
-20.7 | 1.8 | 2.97 | 5.78 | 42.7 | 48.21
-20.6 | 1.8 | 3.08 | 6.04 | 43.26 | 48.48
-20.5 | 1.79 | 3.18 | 6.33 | 43.78 | 48.76
-20.4 | 1.9 | 3.28 | 6.71 | 44.28 | 49.06
-20.3 | 1.9 | 3.38 | 7.27 | 44.79 | 49.4
-20.2 | 2 | 3.48 | 8.09 | 45.26 | 49.78
-20.1 | 2 | 3.59 | 9.49 | 45.65 | 50.08
-20 | 2 | 3.66 | 11.5 | 46.06 | 50.37
-19.9 | 2.1 | 3.85 | 14.2 | 46.46 | 50.73
-19.8 | 2.2 | 3.97 | 17.67 | 46.87 | 51.09
-19.7 | 2.2 | 4.13 | 21.3 | 47.27 | 51.44
-19.6 | 2.3 | 4.42 | 25.58 | 47.68 | 51.8
-19.5 | 2.4 | 4.72 | 30.08 | 48.06 | 52.05
-19.4 | 2.4 | 5.08 | 34.34 | 48.4 | 52.36
-19.3 | 2.49 | 5.71 | 37.98 | 48.68 | 52.7
-19.2 | 2.59 | 6.63 | 40.81 | 49.07 | 53.08
-19.1 | 2.69 | 8.14 | 42.78 | 49.49 | 53.37
-19 | 2.79 | 10.62 | 44.28 | 49.72 | 53.69
-18.9 | 2.99 | 13.52 | 45.35 | 50.16 | 53.95
-18.8 | 3.19 | 17.47 | 46.3 | 50.51 | 54.17
-18.7 | 3.48 | 22.21 | 46.89 | 50.85 | 54.52
-18.6 | 3.98 | 27.35 | 47.46 | 51.23 | 54.94
-18.5 | 4.78 | 32.68 | 47.99 | 51.54 | 55.23
-18.4 | 6.28 | 37.55 | 48.58 | 51.88 | 55.48
-18.3 | 8.7 | 41.59 | 49.04 | 52.21 | 55.85
-18.2 | 11.62 | 44.52 | 49.51 | 52.59 | 56.05
-18.1 | 15.9 | 46.52 | 49.95 | 52.92 | 56.4
-18 | 21.07 | 47.97 | 50.34 | 53.31 | 56.79
-17.9 | 26.98 | 48.87 | 50.83 | 53.6 | 57.07
-17.8 | 33.05 | 49.64 | 51.26 | 53.9 | 57.42
-17.7 | 38.76 | 50.42 | 51.66 | 54.22 | 57.79
-17.6 | 43.7 | 50.97 | 52.04 | 54.56 | 58.04
-17.5 | 47.13 | 51.55 | 52.34 | 54.89 | 58.38
-17.4 | 49.5 | 52.02 | 52.72 | 55.27 | 58.75
-17.3 | 51.05 | 52.49 | 53.06 | 55.59 | 58.95
-17.2 | 52.17 | 52.92 | 53.35 | 55.98 | 59.29
-17.1 | 52.99 | 53.41 | 53.73 | 56.3 | 59.71
-17 | 53.53 | 53.76 | 54.14 | 56.69 | 60.06
-16.9 | 54.02 | 54.14 | 54.54 | 56.98 | 60.27
-16.8 | 54.54 | 54.46 | 54.95 | 57.28 | 60.58
-16.7 | 54.86 | 54.87 | 55.26 | 57.54 | 60.84
-16.6 | 55.25 | 55.24 | 55.52 | 57.87 | 61.15
-16.5 | 55.66 | 55.54 | 55.9 | 58.24 | 61.43
-16.4 | 55.97 | 55.83 | 56.25 | 58.46 | 61.66
-16.3 | 56.37 | 56.21 | 56.53 | 58.87 | 62.07
-16.2 | 56.68 | 56.56 | 56.92 | 59.18 | 62.38
-16.1 | 56.98 | 56.84 | 57.25 | 59.55 | 62.78
-16 | 57.34 | 57.22 | 57.61 | 59.77 | 63.07
-15.9 | 57.59 | 57.55 | 58.05 | 60.21 | 63.37
-15.8 | 57.85 | 57.93 | 58.29 | 60.56 | 63.7
-15.7 | 58.15 | 58.27 | 58.63 | 60.76 | 64.1
-15.6 | 58.45 | 58.57 | 58.98 | 61.11 | 64.36
-15.5 | 58.76 | 58.84 | 59.24 | 61.48 | 64.54
-15.4 | 58.97 | 59.13 | 59.52 | 61.75 | 64.87
-15.3 | 59.27 | 59.43 | 59.9 | 62.16 | 65.27
-15.2 | 59.47 | 59.74 | 60.25 | 62.45 | 65.56
-15.1 | 59.86 | 60.03 | 60.53 | 62.76 | 65.87
-15 | 60.08 | 60.42 | 60.91 | 63.05 | 66.15
-14.9 | 60.37 | 60.68 | 61.26 | 63.38 | 66.36
-14.8 | 60.68 | 61.01 | 61.54 | 63.78 | 66.68
-14.7 | 60.94 | 61.37 | 61.94 | 64.04 | 66.97
-14.6 | 61.24 | 61.63 | 62.19 | 64.28 | 67.24
-14.5 | 61.54 | 61.91 | 62.52 | 64.69 | 67.55
-14.4 | 61.86 | 62.22 | 62.88 | 64.96 | 67.86
-14.3 | 62.06 | 62.52 | 63.14 | 65.24 | 68.26
-14.2 | 62.35 | 62.83 | 63.42 | 65.65 | 68.56
-14.1 | 62.66 | 63.13 | 63.79 | 65.94 | 68.86
-14 | 62.96 | 63.42 | 64.15 | 66.24 | 69.18
-13.9 | 63.26 | 63.8 | 64.44 | 66.56 | 69.58
-13.8 | 63.57 | 64.15 | 64.75 | 66.96 | 69.86
-13.7 | 63.75 | 64.41 | 65.04 | 67.25 | 70.07
-13.6 | 64.02 | 64.71 | 65.41 | 67.58 | 70.34
-13.5 | 64.33 | 64.95 | 65.77 | 67.95 | 70.64
-13.4 | 64.63 | 65.25 | 66.03 | 68.12 | 70.95
-13.3 | 64.94 | 65.62 | 66.3 | 68.53 | 71.26
-13.2 | 65.25 | 65.93 | 66.68 | 68.82 | 71.66
-13.1 | 65.46 | 66.16 | 66.98 | 69.03 | 71.96
-13 | 65.74 | 66.47 | 67.2 | 69.44 | 72.15
-12.9 | 66.05 | 66.85 | 67.61 | 69.74 | 72.38
-12.8 | 66.36 | 67.07 | 67.87 | 70.15 | 72.89
-12.7 | 66.57 | 67.31 | 68.24 | 70.44 | 73.17
-12.6 | 66.85 | 67.7 | 68.67 | 70.74 | 73.48
-12.5 | 67.21 | 68.03 | 68.87 | 71 | 73.74
-12.4 | 67.53 | 68.29 | 69.17 | 71.31 | 74.05
-12.3 | 67.75 | 68.62 | 69.52 | 71.71 | 74.35
-12.2 | 68.02 | 68.85 | 69.81 | 71.91 | 74.55
-12.1 | 68.33 | 69.15 | 70.12 | 72.32 | 74.86
-12 | 68.63 | 69.53 | 70.42 | 72.62 | 75.16
-11.9 | 68.94 | 69.81 | 70.71 | 72.92 | 75.47
-11.8 | 69.24 | 70.12 | 71.07 | 73.23 | 75.88
-11.7 | 69.55 | 70.43 | 71.45 | 73.52 | 76.17
-11.6 | 69.75 | 70.71 | 71.67 | 73.79 | 76.43
-11.5 | 70.13 | 70.98 | 71.95 | 74.19 | 76.74
-11.4 | 70.44 | 71.29 | 72.3 | 74.5 | 76.84
-11.3 | 70.62 | 71.59 | 72.6 | 74.8 | 77.25
-11.2 | 71 | 71.9 | 72.9 | 75 | 77.55
-11.1 | 71.25 | 72.2 | 73.21 | 75.41 | 77.76
-11 | 71.5 | 72.51 | 73.5 | 75.71 | 78.16
-10.9 | 71.92 | 72.81 | 73.86 | 76.01 | 78.46
-10.8 | 72.14 | 73.12 | 74.23 | 76.42 | 78.76
-10.7 | 72.51 | 73.4 | 74.52 | 76.68 | 78.93
-10.6 | 72.84 | 73.78 | 74.81 | 76.98 | 79.33
-10.5 | 73.14 | 74.06 | 75.19 | 77.27 | 79.64
-10.4 | 73.44 | 74.3 | 75.48 | 77.69 | 79.94
-10.3 | 73.75 | 74.7 | 75.67 | 77.79 | 80.25
-10.2 | 74.01 | 74.98 | 76.04 | 78.19 | 80.45
-10.1 | 74.3 | 75.29 | 76.41 | 78.39 | 80.75
-10 | 74.61 | 75.59 | 76.7 | 78.79 | 81.06
-9.9 | 74.91 | 75.9 | 77 | 79 | 81.36
-9.8 | 75.22 | 76.2 | 77.31 | 79.3 | 81.64
-9.7 | 75.51 | 76.52 | 77.61 | 79.66 | 81.82
-9.6 | 75.91 | 76.75 | 77.92 | 79.88 | 82.23
-9.5 | 76.15 | 76.98 | 78.19 | 80.15 | 82.35
-9.4 | 76.41 | 77.43 | 78.5 | 80.66 | 82.72
-9.3 | 76.81 | 77.75 | 78.89 | 80.88 | 83.05
-9.2 | 77.14 | 77.99 | 79.17 | 81.17 | 83.24
-9.1 | 77.4 | 78.44 | 79.52 | 81.49 | 83.65
-9 | 77.69 | 78.76 | 79.91 | 81.68 | 83.86
-8.9 | 78 | 78.98 | 80.16 | 82.08 | 84.16
-8.8 | 78.29 | 79.3 | 80.38 | 82.39 | 84.34
-8.7 | 78.69 | 79.59 | 80.74 | 82.67 | 84.62
-8.6 | 78.93 | 79.92 | 81.12 | 82.85 | 84.93
-8.5 | 79.28 | 80.27 | 81.39 | 83.26 | 85.23
-8.4 | 79.62 | 80.53 | 81.67 | 83.38 | 85.53
-8.3 | 79.92 | 80.78 | 81.97 | 83.75 | 85.84
-8.2 | 80.21 | 81.18 | 82.28 | 84.07 | 86.15
-8.1 | 80.61 | 81.47 | 82.55 | 84.37 | 86.35
-8 | 80.91 | 81.78 | 82.75 | 84.57 | 86.65
-7.9 | 81.18 | 82.07 | 83.16 | 84.97 | 86.96
-7.8 | 81.49 | 82.43 | 83.58 | 85.23 | 87.13
-7.7 | 81.8 | 82.75 | 83.77 | 85.63 | 87.42
-7.6 | 82.09 | 83.01 | 84.14 | 85.86 | 87.64
-7.5 | 82.47 | 83.39 | 84.47 | 86.14 | 87.92
-7.4 | 82.81 | 83.66 | 84.64 | 86.46 | 88.24
-7.3 | 83.1 | 83.92 | 84.99 | 86.66 | 88.45
-7.2 | 83.41 | 84.2 | 85.38 | 86.95 | 88.65
-7.1 | 83.72 | 84.66 | 85.66 | 87.26 | 88.94
-7 | 84 | 84.85 | 85.98 | 87.56 | 89.24
-6.9 | 84.4 | 85.17 | 86.24 | 87.86 | 89.55
-6.8 | 84.7 | 85.5 | 86.5 | 88.13 | 89.75
-6.7 | 84.98 | 85.9 | 86.9 | 88.43 | 89.92
-6.6 | 85.29 | 86.13 | 87.18 | 88.65 | 90.3
-6.5 | 85.59 | 86.35 | 87.49 | 88.94 | 90.55
-6.4 | 85.89 | 86.75 | 87.73 | 89.14 | 90.74
-6.3 | 86.19 | 87.05 | 87.94 | 89.52 | 90.95
-6.2 | 86.5 | 87.34 | 88.28 | 89.76 | 91.14
-6.1 | 86.79 | 87.68 | 88.67 | 90.04 | 91.44
-6 | 87.17 | 88.04 | 88.94 | 90.35 | 91.65
-5.9 | 87.52 | 88.28 | 89.16 | 90.57 | 91.95
-5.8 | 87.8 | 88.65 | 89.47 | 90.75 | 92.13
-5.7 | 88.17 | 88.86 | 89.77 | 90.92 | 92.42
-5.6 | 88.41 | 89.17 | 90.07 | 91.29 | 92.57
-5.5 | 88.72 | 89.42 | 90.39 | 91.56 | 92.74
-5.4 | 89.1 | 89.75 | 90.64 | 91.74 | 92.94
-5.3 | 89.31 | 90.15 | 90.85 | 92.02 | 93.21
-5.2 | 89.64 | 90.44 | 91.24 | 92.33 | 93.61
-5.1 | 90 | 90.75 | 91.44 | 92.64 | 93.88
-5 | 90.29 | 91.03 | 91.75 | 92.86 | 93.96
-4.9 | 90.59 | 91.26 | 92.06 | 93.06 | 94.24
-4.8 | 90.9 | 91.65 | 92.33 | 93.23 | 94.46
-4.7 | 91.19 | 91.84 | 92.58 | 93.49 | 94.62
-4.6 | 91.59 | 92.17 | 92.93 | 93.8 | 94.92
-4.5 | 91.83 | 92.46 | 93.06 | 94.13 | 95.08
-4.4 | 92.05 | 92.72 | 93.43 | 94.26 | 95.24
-4.3 | 92.37 | 93.03 | 93.61 | 94.51 | 95.45
-4.2 | 92.68 | 93.32 | 93.83 | 94.75 | 95.65
-4.1 | 92.91 | 93.53 | 94.14 | 94.93 | 95.85
-4 | 93.23 | 93.84 | 94.43 | 95.22 | 96.07
-3.9 | 93.61 | 94.13 | 94.64 | 95.47 | 96.16
-3.8 | 93.8 | 94.34 | 94.96 | 95.56 | 96.39
-3.7 | 94.16 | 94.63 | 95.23 | 95.83 | 96.61
-3.6 | 94.43 | 94.84 | 95.32 | 96.06 | 96.58
-3.5 | 94.69 | 95.11 | 95.52 | 96.16 | 96.81
-3.4 | 94.93 | 95.31 | 95.82 | 96.31 | 96.99
-3.3 | 95.18 | 95.61 | 96.01 | 96.6 | 97.14
-3.2 | 95.4 | 95.82 | 96.24 | 96.85 | 97.35
-3.1 | 95.66 | 96.11 | 96.61 | 96.96 | 97.56
-3 | 95.92 | 96.21 | 96.6 | 97.15 | 97.69
-2.9 | 96.1 | 96.54 | 96.83 | 97.26 | 97.78
-2.8 | 96.37 | 96.8 | 97.01 | 97.51 | 97.95
-2.7 | 96.63 | 96.81 | 97.13 | 97.6 | 98.07
-2.6 | 96.82 | 97.1 | 97.44 | 97.72 | 98.22
-2.5 | 97.02 | 97.4 | 97.63 | 97.79 | 98.46
-2.4 | 97.21 | 97.6 | 97.81 | 97.97 | 98.52
-2.3 | 97.5 | 97.71 | 98.01 | 98.25 | 98.57
-2.2 | 97.67 | 98 | 98.1 | 98.37 | 98.67
-2.1 | 97.78 | 98 | 98.22 | 98.46 | 98.85
-2 | 98.02 | 98.32 | 98.52 | 98.57 | 98.98
-1.9 | 98.14 | 98.5 | 98.61 | 98.66 | 99.16
-1.8 | 98.36 | 98.5 | 98.71 | 98.75 | 99.17
-1.7 | 98.56 | 98.72 | 98.82 | 98.86 | 99.15
-1.6 | 98.71 | 98.92 | 99.01 | 98.94 | 99.28
-1.5 | 98.87 | 99.09 | 99.01 | 99.12 | 99.37
-1.4 | 98.94 | 99.19 | 99.22 | 99.26 | 99.49
-1.3 | 99.12 | 99.29 | 99.3 | 99.38 | 99.51
-1.2 | 99.27 | 99.4 | 99.4 | 99.33 | 99.63
-1.1 | 99.37 | 99.5 | 99.5 | 99.41 | 99.8
-1 | 99.39 | 99.6 | 99.6 | 99.56 | 99.83
-0.9 | 99.48 | 99.69 | 99.61 | 99.66 | 99.81
-0.8 | 99.67 | 99.7 | 99.81 | 99.65 | 99.87
-0.7 | 99.69 | 99.9 | 99.8 | 99.64 | 99.93
-0.6 | 99.75 | 99.9 | 99.81 | 99.7 | 99.92
-0.5 | 99.81 | 100 | 99.9 | 99.75 | 99.92
-0.4 | 99.79 | 99.99 | 99.9 | 99.8 | 99.92
-0.3 | 99.86 | 100 | 100 | 99.87 | 99.91
-0.2 | 99.9 | 99.89 | 100 | 99.84 | 99.98
-0.1 | 99.93 | 99.99 | 100 | 99.86 | 100
0 | 100 | 100 | 100 | 100 | 99.92
0.1 | 99.93 | 99.91 | 99.99 | 99.86 | 100
0.2 | 99.9 | 99.9 | 100 | 99.84 | 99.98
0.3 | 99.86 | 100 | 100 | 99.87 | 99.91
0.4 | 99.79 | 99.99 | 99.9 | 99.8 | 99.92
0.5 | 99.81 | 100 | 99.9 | 99.75 | 99.92
0.6 | 99.75 | 99.9 | 99.81 | 99.7 | 99.92
0.7 | 99.69 | 99.9 | 99.8 | 99.64 | 99.93
0.8 | 99.67 | 99.7 | 99.81 | 99.65 | 99.87
0.9 | 99.48 | 99.69 | 99.61 | 99.66 | 99.81
1 | 99.39 | 99.6 | 99.6 | 99.56 | 99.83
1.1 | 99.37 | 99.5 | 99.5 | 99.41 | 99.8
1.2 | 99.27 | 99.4 | 99.4 | 99.33 | 99.63
1.3 | 99.12 | 99.29 | 99.3 | 99.38 | 99.51
1.4 | 98.94 | 99.19 | 99.22 | 99.26 | 99.49
1.5 | 98.87 | 99.09 | 99.01 | 99.12 | 99.37
1.6 | 98.71 | 98.92 | 99.01 | 98.94 | 99.28
1.7 | 98.56 | 98.72 | 98.82 | 98.86 | 99.15
1.8 | 98.36 | 98.5 | 98.71 | 98.75 | 99.17
1.9 | 98.14 | 98.5 | 98.61 | 98.66 | 99.16
2 | 98.02 | 98.32 | 98.52 | 98.57 | 98.98
2.1 | 97.78 | 98 | 98.22 | 98.46 | 98.85
2.2 | 97.67 | 98 | 98.1 | 98.37 | 98.67
2.3 | 97.5 | 97.71 | 98.01 | 98.25 | 98.57
2.4 | 97.21 | 97.6 | 97.81 | 97.97 | 98.52
2.5 | 97.02 | 97.4 | 97.63 | 97.79 | 98.46
2.6 | 96.82 | 97.1 | 97.44 | 97.72 | 98.22
2.7 | 96.63 | 96.81 | 97.13 | 97.6 | 98.07
2.8 | 96.37 | 96.8 | 97.01 | 97.51 | 97.95
2.9 | 96.1 | 96.54 | 96.83 | 97.26 | 97.78
3 | 95.92 | 96.21 | 96.6 | 97.15 | 97.69
3.1 | 95.66 | 96.11 | 96.61 | 96.96 | 97.56
3.2 | 95.4 | 95.82 | 96.24 | 96.85 | 97.35
3.3 | 95.18 | 95.61 | 96.01 | 96.6 | 97.14
3.4 | 94.93 | 95.31 | 95.82 | 96.31 | 96.99
3.5 | 94.69 | 95.11 | 95.52 | 96.16 | 96.81
3.6 | 94.43 | 94.84 | 95.32 | 96.06 | 96.58
3.7 | 94.16 | 94.63 | 95.23 | 95.83 | 96.61
3.8 | 93.8 | 94.34 | 94.96 | 95.56 | 96.39
3.9 | 93.61 | 94.13 | 94.64 | 95.47 | 96.16
4 | 93.23 | 93.84 | 94.43 | 95.22 | 96.07
4.1 | 92.91 | 93.53 | 94.14 | 94.93 | 95.85
4.2 | 92.68 | 93.32 | 93.83 | 94.75 | 95.65
4.3 | 92.37 | 93.03 | 93.61 | 94.51 | 95.45
4.4 | 92.05 | 92.72 | 93.43 | 94.26 | 95.24
4.5 | 91.83 | 92.46 | 93.06 | 94.13 | 95.08
4.6 | 91.59 | 92.17 | 92.93 | 93.8 | 94.92
4.7 | 91.19 | 91.84 | 92.58 | 93.49 | 94.62
4.8 | 90.9 | 91.65 | 92.33 | 93.23 | 94.46
4.9 | 90.59 | 91.26 | 92.06 | 93.06 | 94.24
5 | 90.29 | 91.03 | 91.75 | 92.86 | 93.96
5.1 | 90 | 90.75 | 91.44 | 92.64 | 93.88
5.2 | 89.64 | 90.44 | 91.24 | 92.33 | 93.61
5.3 | 89.31 | 90.15 | 90.85 | 92.02 | 93.21
5.4 | 89.1 | 89.75 | 90.64 | 91.74 | 92.94
5.5 | 88.72 | 89.42 | 90.39 | 91.56 | 92.74
5.6 | 88.41 | 89.17 | 90.07 | 91.29 | 92.57
5.7 | 88.17 | 88.86 | 89.77 | 90.92 | 92.42
5.8 | 87.8 | 88.65 | 89.47 | 90.75 | 92.13
5.9 | 87.52 | 88.28 | 89.16 | 90.57 | 91.95
6 | 87.17 | 88.04 | 88.94 | 90.35 | 91.65
6.1 | 86.79 | 87.68 | 88.67 | 90.04 | 91.44
6.2 | 86.5 | 87.34 | 88.28 | 89.76 | 91.14
6.3 | 86.19 | 87.05 | 87.94 | 89.52 | 90.95
6.4 | 85.89 | 86.75 | 87.73 | 89.14 | 90.74
6.5 | 85.59 | 86.35 | 87.49 | 88.94 | 90.55
6.6 | 85.29 | 86.13 | 87.18 | 88.65 | 90.3
6.7 | 84.98 | 85.9 | 86.9 | 88.43 | 89.92
6.8 | 84.7 | 85.5 | 86.5 | 88.13 | 89.75
6.9 | 84.4 | 85.17 | 86.24 | 87.86 | 89.55
7 | 84 | 84.85 | 85.98 | 87.56 | 89.24
7.1 | 83.72 | 84.66 | 85.66 | 87.26 | 88.94
7.2 | 83.41 | 84.2 | 85.38 | 86.95 | 88.65
7.3 | 83.1 | 83.92 | 84.99 | 86.66 | 88.45
7.4 | 82.81 | 83.66 | 84.64 | 86.46 | 88.24
7.5 | 82.47 | 83.39 | 84.47 | 86.14 | 87.92
7.6 | 82.09 | 83.01 | 84.14 | 85.86 | 87.64
7.7 | 81.8 | 82.75 | 83.77 | 85.63 | 87.42
7.8 | 81.49 | 82.43 | 83.58 | 85.23 | 87.13
7.9 | 81.18 | 82.07 | 83.16 | 84.97 | 86.96
8 | 80.91 | 81.78 | 82.75 | 84.57 | 86.65
8.1 | 80.61 | 81.47 | 82.55 | 84.37 | 86.35
8.2 | 80.21 | 81.18 | 82.28 | 84.07 | 86.15
8.3 | 79.92 | 80.78 | 81.97 | 83.75 | 85.84
8.4 | 79.62 | 80.53 | 81.67 | 83.38 | 85.53
8.5 | 79.28 | 80.27 | 81.39 | 83.26 | 85.23
8.6 | 78.93 | 79.92 | 81.12 | 82.85 | 84.93
8.7 | 78.69 | 79.59 | 80.74 | 82.67 | 84.62
8.8 | 78.29 | 79.3 | 80.38 | 82.39 | 84.34
8.9 | 78 | 78.98 | 80.16 | 82.08 | 84.16
9 | 77.69 | 78.76 | 79.91 | 81.68 | 83.86
9.1 | 77.4 | 78.44 | 79.52 | 81.49 | 83.65
9.2 | 77.14 | 77.99 | 79.17 | 81.17 | 83.24
9.3 | 76.81 | 77.75 | 78.89 | 80.88 | 83.05
9.4 | 76.41 | 77.43 | 78.5 | 80.66 | 82.72
9.5 | 76.15 | 76.98 | 78.19 | 80.15 | 82.35
9.6 | 75.91 | 76.75 | 77.92 | 79.88 | 82.23
9.7 | 75.51 | 76.52 | 77.61 | 79.66 | 81.82
9.8 | 75.22 | 76.2 | 77.31 | 79.3 | 81.64
9.9 | 74.91 | 75.9 | 77 | 79 | 81.36
10 | 74.61 | 75.59 | 76.7 | 78.79 | 81.06
10.1 | 74.3 | 75.29 | 76.41 | 78.39 | 80.75
10.2 | 74.01 | 74.98 | 76.04 | 78.19 | 80.45
10.3 | 73.75 | 74.7 | 75.67 | 77.79 | 80.25
10.4 | 73.44 | 74.3 | 75.48 | 77.69 | 79.94
10.5 | 73.14 | 74.06 | 75.19 | 77.27 | 79.64
10.6 | 72.84 | 73.78 | 74.81 | 76.98 | 79.33
10.7 | 72.51 | 73.4 | 74.52 | 76.68 | 78.93
10.8 | 72.14 | 73.12 | 74.23 | 76.42 | 78.76
10.9 | 71.92 | 72.81 | 73.86 | 76.01 | 78.46
11 | 71.5 | 72.51 | 73.5 | 75.71 | 78.16
11.1 | 71.25 | 72.2 | 73.21 | 75.41 | 77.76
11.2 | 71 | 71.9 | 72.9 | 75 | 77.55
11.3 | 70.62 | 71.59 | 72.6 | 74.8 | 77.25
11.4 | 70.44 | 71.29 | 72.3 | 74.5 | 76.84
11.5 | 70.13 | 70.98 | 71.95 | 74.19 | 76.74
11.6 | 69.75 | 70.71 | 71.67 | 73.79 | 76.43
11.7 | 69.55 | 70.43 | 71.45 | 73.52 | 76.17
11.8 | 69.24 | 70.12 | 71.07 | 73.23 | 75.88
11.9 | 68.94 | 69.81 | 70.71 | 72.92 | 75.47
12 | 68.63 | 69.53 | 70.42 | 72.62 | 75.16
12.1 | 68.33 | 69.15 | 70.12 | 72.32 | 74.86
12.2 | 68.02 | 68.85 | 69.81 | 71.91 | 74.55
12.3 | 67.75 | 68.62 | 69.52 | 71.71 | 74.35
12.4 | 67.53 | 68.29 | 69.17 | 71.31 | 74.05
12.5 | 67.21 | 68.03 | 68.87 | 71 | 73.74
12.6 | 66.85 | 67.7 | 68.67 | 70.74 | 73.48
12.7 | 66.57 | 67.31 | 68.24 | 70.44 | 73.17
12.8 | 66.36 | 67.07 | 67.87 | 70.15 | 72.89
12.9 | 66.05 | 66.85 | 67.61 | 69.74 | 72.38
13 | 65.74 | 66.47 | 67.2 | 69.44 | 72.15
13.1 | 65.46 | 66.16 | 66.98 | 69.03 | 71.96
13.2 | 65.25 | 65.93 | 66.68 | 68.82 | 71.66
13.3 | 64.94 | 65.62 | 66.3 | 68.53 | 71.26
13.4 | 64.63 | 65.25 | 66.03 | 68.12 | 70.95
13.5 | 64.33 | 64.95 | 65.77 | 67.95 | 70.64
13.6 | 64.02 | 64.71 | 65.41 | 67.58 | 70.34
13.7 | 63.75 | 64.41 | 65.04 | 67.25 | 70.07
13.8 | 63.57 | 64.15 | 64.75 | 66.96 | 69.86
13.9 | 63.26 | 63.8 | 64.44 | 66.56 | 69.58
14 | 62.96 | 63.42 | 64.15 | 66.24 | 69.18
14.1 | 62.66 | 63.13 | 63.79 | 65.94 | 68.86
14.2 | 62.35 | 62.83 | 63.42 | 65.65 | 68.56
14.3 | 62.06 | 62.52 | 63.14 | 65.24 | 68.26
14.4 | 61.86 | 62.22 | 62.88 | 64.96 | 67.86
14.5 | 61.54 | 61.91 | 62.52 | 64.69 | 67.55
14.6 | 61.24 | 61.63 | 62.19 | 64.28 | 67.24
14.7 | 60.94 | 61.37 | 61.94 | 64.04 | 66.97
14.8 | 60.68 | 61.01 | 61.54 | 63.78 | 66.68
14.9 | 60.37 | 60.68 | 61.26 | 63.38 | 66.36
15 | 60.08 | 60.42 | 60.91 | 63.05 | 66.15
15.1 | 59.86 | 60.03 | 60.53 | 62.76 | 65.87
15.2 | 59.47 | 59.74 | 60.25 | 62.45 | 65.56
15.3 | 59.27 | 59.43 | 59.9 | 62.16 | 65.27
15.4 | 58.97 | 59.13 | 59.52 | 61.75 | 64.87
15.5 | 58.76 | 58.84 | 59.24 | 61.48 | 64.54
15.6 | 58.45 | 58.57 | 58.98 | 61.11 | 64.36
15.7 | 58.15 | 58.27 | 58.63 | 60.76 | 64.1
15.8 | 57.85 | 57.93 | 58.29 | 60.56 | 63.7
15.9 | 57.59 | 57.55 | 58.05 | 60.21 | 63.37
16 | 57.34 | 57.22 | 57.61 | 59.77 | 63.07
16.1 | 56.98 | 56.84 | 57.25 | 59.55 | 62.78
16.2 | 56.68 | 56.56 | 56.92 | 59.18 | 62.38
16.3 | 56.37 | 56.21 | 56.53 | 58.87 | 62.07
16.4 | 55.97 | 55.83 | 56.25 | 58.46 | 61.66
16.5 | 55.66 | 55.54 | 55.9 | 58.24 | 61.43
16.6 | 55.25 | 55.24 | 55.52 | 57.87 | 61.15
16.7 | 54.86 | 54.87 | 55.26 | 57.54 | 60.84
16.8 | 54.54 | 54.46 | 54.95 | 57.28 | 60.58
16.9 | 54.02 | 54.14 | 54.54 | 56.98 | 60.27
17 | 53.53 | 53.76 | 54.14 | 56.69 | 60.06
17.1 | 52.99 | 53.41 | 53.73 | 56.3 | 59.71
17.2 | 52.17 | 52.92 | 53.35 | 55.98 | 59.29
17.3 | 51.05 | 52.49 | 53.06 | 55.59 | 58.95
17.4 | 49.5 | 52.02 | 52.72 | 55.27 | 58.75
17.5 | 47.13 | 51.55 | 52.34 | 54.89 | 58.38
17.6 | 43.7 | 50.97 | 52.04 | 54.56 | 58.04
17.7 | 38.76 | 50.42 | 51.66 | 54.22 | 57.79
17.8 | 33.05 | 49.64 | 51.26 | 53.9 | 57.42
17.9 | 26.98 | 48.87 | 50.83 | 53.6 | 57.07
18 | 21.07 | 47.97 | 50.34 | 53.31 | 56.79
18.1 | 15.9 | 46.52 | 49.95 | 52.92 | 56.4
18.2 | 11.62 | 44.52 | 49.51 | 52.59 | 56.05
18.3 | 8.7 | 41.59 | 49.04 | 52.21 | 55.85
18.4 | 6.28 | 37.55 | 48.58 | 51.88 | 55.48
18.5 | 4.78 | 32.68 | 47.99 | 51.54 | 55.23
18.6 | 3.98 | 27.35 | 47.46 | 51.23 | 54.94
18.7 | 3.48 | 22.21 | 46.89 | 50.85 | 54.52
18.8 | 3.19 | 17.47 | 46.3 | 50.51 | 54.17
18.9 | 2.99 | 13.52 | 45.35 | 50.16 | 53.95
19 | 2.79 | 10.62 | 44.28 | 49.72 | 53.69
19.1 | 2.69 | 8.14 | 42.78 | 49.49 | 53.37
19.2 | 2.59 | 6.63 | 40.81 | 49.07 | 53.08
19.3 | 2.49 | 5.71 | 37.98 | 48.68 | 52.7
19.4 | 2.4 | 5.08 | 34.34 | 48.4 | 52.36
19.5 | 2.4 | 4.72 | 30.08 | 48.06 | 52.05
19.6 | 2.3 | 4.42 | 25.58 | 47.68 | 51.8
19.7 | 2.2 | 4.13 | 21.3 | 47.27 | 51.44
19.8 | 2.2 | 3.97 | 17.67 | 46.87 | 51.09
19.9 | 2.1 | 3.85 | 14.2 | 46.46 | 50.73
20 | 2 | 3.66 | 11.5 | 46.06 | 50.37
20.1 | 2 | 3.59 | 9.49 | 45.65 | 50.08
20.2 | 2 | 3.48 | 8.09 | 45.26 | 49.78
20.3 | 1.9 | 3.38 | 7.27 | 44.79 | 49.4
20.4 | 1.9 | 3.28 | 6.71 | 44.28 | 49.06
20.5 | 1.79 | 3.18 | 6.33 | 43.78 | 48.76
20.6 | 1.8 | 3.08 | 6.04 | 43.26 | 48.48
20.7 | 1.8 | 2.97 | 5.78 | 42.7 | 48.21
20.8 | 1.7 | 2.9 | 5.58 | 41.99 | 47.85
20.9 | 1.7 | 2.89 | 5.39 | 41.12 | 47.42
21 | 1.7 | 2.77 | 5.28 | 40.1 | 47.08
21.1 | 1.6 | 2.7 | 5.08 | 38.52 | 46.78
21.2 | 1.6 | 2.69 | 4.99 | 36.55 | 46.49
21.3 | 1.6 | 2.6 | 4.88 | 34 | 46.11
21.4 | 1.5 | 2.59 | 4.68 | 31 | 45.76
21.5 | 1.5 | 2.48 | 4.59 | 27.49 | 45.47
21.6 | 1.5 | 2.4 | 4.5 | 24.11 | 45.14
21.7 | 1.5 | 2.39 | 4.4 | 20.89 | 44.75
21.8 | 1.41 | 2.3 | 4.3 | 17.69 | 44.33
21.9 | 1.4 | 2.29 | 4.2 | 15.01 | 43.99
22 | 1.4 | 2.2 | 4.1 | 12.89 | 43.64
22.1 | 1.4 | 2.2 | 3.99 | 11.36 | 43.22
22.2 | 1.4 | 2.1 | 3.89 | 10.27 | 42.82
22.3 | 1.3 | 2.11 | 3.79 | 9.55 | 42.41
22.4 | 1.3 | 2.09 | 3.7 | 9 | 42.03
22.5 | 1.3 | 1.99 | 3.7 | 8.61 | 41.51
22.6 | 1.3 | 2.01 | 3.59 | 8.34 | 40.86
22.7 | 1.3 | 1.99 | 3.49 | 8.11 | 40.28
22.8 | 1.21 | 1.9 | 3.4 | 7.9 | 39.54
22.9 | 1.2 | 1.89 | 3.4 | 7.75 | 38.57
23 | 1.2 | 1.79 | 3.3 | 7.61 | 37.44
23.1 | 1.2 | 1.8 | 3.3 | 7.4 | 35.95
23.2 | 1.2 | 1.8 | 3.19 | 7.19 | 33.96
23.3 | 1.2 | 1.8 | 3.1 | 7.05 | 31.57
23.4 | 1.1 | 1.7 | 3.1 | 6.9 | 28.86
23.5 | 1.1 | 1.7 | 3 | 6.74 | 25.97
23.6 | 1.1 | 1.7 | 3 | 6.67 | 22.93
23.7 | 1.1 | 1.69 | 2.9 | 6.56 | 20.15
23.8 | 1.1 | 1.59 | 2.9 | 6.47 | 17.82
23.9 | 1.1 | 1.61 | 2.8 | 6.32 | 15.69
24 | 1.01 | 1.59 | 2.8 | 6.15 | 13.94
24.1 | 1 | 1.49 | 2.7 | 6.06 | 12.59
24.2 | 1 | 1.5 | 2.7 | 5.95 | 11.69
24.3 | 1 | 1.5 | 2.6 | 5.85 | 11.04
24.4 | 1 | 1.5 | 2.5 | 5.75 | 10.52
24.5 | 1 | 1.5 | 2.5 | 5.66 | 10.18
24.6 | 1 | 1.4 | 2.5 | 5.57 | 9.92
24.7 | 0.9 | 1.4 | 2.4 | 5.46 | 9.76
24.8 | 0.9 | 1.4 | 2.4 | 5.36 | 9.51
24.9 | 0.9 | 1.4 | 2.41 | 5.26 | 9.28
25 | 0.9 | 1.39 | 2.3 | 5.15 | 9.16
25.1 | 0.9 | 1.29 | 2.3 | 5.1 | 8.99
25.2 | 0.9 | 1.3 | 2.3 | 5.07 | 8.85
25.3 | 0.9 | 1.3 | 2.2 | 4.96 | 8.68
25.4 | 0.9 | 1.3 | 2.2 | 4.85 | 8.55
25.5 | 0.9 | 1.3 | 2.1 | 4.8 | 8.37
25.6 | 0.91 | 1.3 | 2.1 | 4.78 | 8.29
25.7 | 0.81 | 1.3 | 2.1 | 4.67 | 8.14
25.8 | 0.8 | 1.2 | 2.01 | 4.56 | 7.99
25.9 | 0.8 | 1.2 | 2 | 4.5 | 7.89
26 | 0.8 | 1.2 | 2 | 4.47 | 7.8
26.1 | 0.8 | 1.2 | 2.01 | 4.36 | 7.65
26.2 | 0.8 | 1.2 | 1.91 | 4.26 | 7.48
26.3 | 0.8 | 1.1 | 1.9 | 4.2 | 7.39
26.4 | 0.8 | 1.1 | 1.9 | 4.17 | 7.28
26.5 | 0.8 | 1.1 | 1.9 | 4.1 | 7.18
26.6 | 0.8 | 1.1 | 1.8 | 4.08 | 7.08
26.7 | 0.8 | 1.1 | 1.8 | 3.96 | 6.99
26.8 | 0.71 | 1.1 | 1.8 | 3.89 | 6.87
26.9 | 0.7 | 1.1 | 1.81 | 3.91 | 6.68
27 | 0.7 | 1 | 1.71 | 3.88 | 6.59
27.1 | 0.7 | 1 | 1.7 | 3.76 | 6.49
27.2 | 0.7 | 1 | 1.71 | 3.7 | 6.38
27.3 | 0.7 | 1 | 1.61 | 3.67 | 6.32
27.4 | 0.7 | 1 | 1.6 | 3.6 | 6.3
27.5 | 0.7 | 1 | 1.6 | 3.58 | 6.18
27.6 | 0.7 | 1 | 1.6 | 3.49 | 6.08
27.7 | 0.7 | 1 | 1.6 | 3.51 | 5.98
27.8 | 0.7 | 1 | 1.51 | 3.48 | 5.89
27.9 |  | 0.9 | 1.5 | 3.37 | 5.79
28 |  | 0.9 | 1.5 | 3.3 | 5.71
28.1 |  | 0.9 | 1.5 | 3.28 | 5.7
28.2 |  | 0.9 | 1.5 | 3.19 | 5.58
28.3 |  | 0.9 | 1.51 | 3.21 | 5.51
28.4 |  | 0.9 | 1.41 | 3.17 | 5.5
28.5 |  | 0.9 | 1.4 | 3.1 | 5.38
28.6 |  |  | 1.4 | 3.09 | 5.32
28.7 |  |  | 1.4 | 2.99 | 5.29
28.8 |  |  | 1.4 | 3.01 | 5.18
28.9 |  |  | 1.41 | 2.98 | 5.08
29 |  |  | 1.32 | 2.9 | 5.01
29.1 |  |  | 1.3 | 2.88 | 5
29.2 |  |  | 1.3 | 2.8 | 4.88
29.3 |  |  | 1.3 | 2.77 | 4.81
29.4 |  |  | 1.3 | 2.71 | 4.79
29.5 |  |  | 1.3 | 2.8 | 4.71
29.6 |  |  |  | 2.69 | 4.7
29.7 |  |  |  | 2.69 | 4.58
29.8 |  |  |  | 2.59 | 4.48
29.9 |  |  |  | 2.6 | 4.41
30 |  |  |  | 2.6 | 4.42
30.1 |  |  |  | 2.58 | 4.4
30.2 |  |  |  | 2.49 | 4.31
30.3 |  |  |  | 2.51 | 4.3
30.4 |  |  |  | 2.49 | 4.18
30.5 |  |  |  | 2.4 | 4.11
30.6 |  |  |  | 2.39 | 4.12
30.7 |  |  |  | 2.29 | 4.1
30.8 |  |  |  | 2.3 | 3.98
30.9 |  |  |  | 2.3 | 3.9
31 |  |  |  | 2.3 | 3.92
31.1 |  |  |  | 2.29 | 3.89
31.2 |  |  |  | 2.19 | 3.81
31.3 |  |  |  | 2.2 | 3.79
31.4 |  |  |  | 2.19 | 3.59
31.5 |  |  |  | 2.08 | 3.62
31.6 |  |  |  |  | 3.61
31.7 |  |  |  |  | 3.62
31.8 |  |  |  |  | 3.59
31.9 |  |  |  |  | 3.5
32 |  |  |  |  | 3.51
32.1 |  |  |  |  | 3.51
32.2
32.3
32.4
32.5
32.6
32.7
32.8
32.9
33
33.1
33.2
33.3
33.4
33.5
33.6
33.7
33.8
33.9
34
34.1
34.2
34.3
34.4
34.5
34.6
34.7
34.8
34.9
35
35.1
35.2
35.3
35.4
35.5
35.6
35.7
35.8
35.9
36
36.1
36.2
36.3
36.4
36.5
36.6
36.7
36.8
36.9
37
37.1
37.2
37.3
37.4
37.5
37.6
37.7
37.8
37.9
38
38.1
38.2
38.3
38.4
38.5
38.6
38.7
38.8
38.9
39
39.1
39.2
39.3
39.4
39.5
39.6
39.7
39.8
39.9
40
40.1
40.2
40.3
40.4
40.5
40.6
40.7
40.8
40.9
41
41.1
41.2
41.3
41.4
41.5
41.6
41.7
41.8
41.9
42
42.1
42.2
42.3
42.4
42.5
42.6
42.7
42.8
42.9
43
43.1
43.2
43.3
43.4
43.5
43.6
43.7
43.8
43.9
44
44.1
44.2
44.3
44.4
44.5
44.6
44.7
44.8
44.9
45
45.1
45.2
45.3
45.4
45.5
45.6
45.7
45.8
45.9
46
46.1
46.2
46.3
46.4
46.5
46.6
46.7
46.8
46.9
47
47.1
47.2
47.3
47.4
47.5
47.6
47.7
47.8
47.9
48
48.1
48.2
48.3
48.4
48.5
48.6
48.7
48.8
48.9
49
49.1
49.2
49.3
49.4
49.5
49.6
49.7
49.8
49.9
50
50.1
50.2
50.3
50.4
50.5
50.6
50.7
50.8
50.9
51
51.1
51.2
51.3
51.4
51.5
51.6
51.7
51.8
51.9
52
52.1
52.2
52.3
52.4
52.5
52.6
52.7
52.8
52.9
53
53.1
53.2
53.3
53.4
53.5
53.6
53.7
53.8
53.9
54
54.1
54.2
54.3
54.4
54.5
54.6
54.7
54.8
54.9
55
55.1
55.2

## Open Field Output Factors
Output factors
Open field
SSD = 95 cm
Depth = 5 cm
Detector: SNC Edge Detector used for 1xX and 1xY, CC13 ionization chamber used for remainder.
Field size X [cm]
Field Size Y [cm] |  | 1 | 2 | 4 | 6 | 8 | 10 | 14 | 20 | 24 | 28
| 1 | 0.766 | 0.792 | 0.804 | 0.808 | 0.81 | 0.811 | 0.813 | 0.814 | 0.814 | 0.814
| 2 | 0.816 | 0.855 | 0.884 | 0.892 | 0.897 | 0.899 | 0.902 | 0.904 | 0.905 | 0.905
| 4 | 0.829 | 0.885 | 0.922 | 0.939 | 0.947 | 0.953 | 0.958 | 0.962 | 0.964 | 0.964
| 6 | 0.834 | 0.893 | 0.939 | 0.958 | 0.971 | 0.977 | 0.985 | 0.991 | 0.994 | 0.994
| 8 | 0.836 | 0.898 | 0.949 | 0.972 | 0.982 | 0.992 | 1.002 | 1.01 | 1.013 | 1.013
| 10 | 0.837 | 0.899 | 0.952 | 0.977 | 0.991 | 1 | 1.012 | 1.021 | 1.024 | 1.024
| 14 | 0.839 | 0.902 | 0.957 | 0.984 | 1.001 | 1.012 | 1.027 | 1.038 | 1.043 | 1.043
| 20 | 0.84 | 0.904 | 0.961 | 0.99 | 1.008 | 1.02 | 1.038 | 1.051 | 1.057 | 1.058
| 24 | 0.84 | 0.905 | 0.962 | 0.992 | 1.011 | 1.024 | 1.042 | 1.057 | 1.062 | 1.064
| 28 | 0.841 | 0.905 | 0.963 | 0.992 | 1.011 | 1.024 | 1.042 | 1.058 | 1.064 | 1.067